## Cover
Main Menu
LT500 Open Channel Flow
LT500 Pump Control
LT500 Differential
Hydroranger-200 Open Channel flow application (except contracted weirs)
Hydroranger-200 HMI Open Channel flow applications (except contracted weirs)
Hydroranger-200 Open Channel flow, contracted weirs
Hydroranger-200 HMI Open Channel flow, contracted weirs
LUT-400 Open Channel flow, (except contracted weirs)
LUT-400 Open Channel flow, contracted weirs
Probe-LU, Probe-LR, Open Channel flow (except contracted weirs)
LR-200, LR-250 Open Channel Flow (Except contracted weirs)
OCM-III Open Channel Flow
Hydroranger-200, Well Pump Control
Hydroranger-200, Travelling Screen
Hydroranger-200 HMI Differential
Keystroke Instructions

## LT500 OCF
LT500 Open Channel Flowmeter
Volume Flow Units
Date: | 2026-07-29
Location:
Primary Device: | 6" Parshall Flume | Sensor Type
Maximum flow rate (MGD)
Go to: Wizards/Quick Start/Volume Flow
Description | Value
Length Unit | Inches | Response Rate
Volume Flow Units | Mgal/d
Volume Units | Custom
Sensor Type | LU240
Lower Calibration Point |  | Primary Measuring Device
Upper calibration point | 0
Response Rate (slow, med, fast) | Fast
Material Type | Liquid
Primary Measuring Device | Exponential
Method of Flow Calculation | Ratiometric
Flow Exponent | 1.58
Maximum head | 0 | Totalizer Units
Zero Head Offset
Maximum Flow | 0
Low Flow cutoff (flow units)
Go to: Setup/Totalizers
Process Value | Volume Flow (pt. 1) | Totalizer Decimal | View 2 (enabled/disabled)
Units | Custom
Custom Units | Mgal
Custom Conversion Factor | 0.0002642
Decimal Places | 2
|  | Display View 1 | Display View 2
Go to: Setup/Local Display
View 1
Type | 3 values, icons and vertical bargraph
1st Value | Volume flow (point 1)
2nd Value | Head (point 1) | View 1, 1st value | View 2, 1st value
3rd Value | Totalizer 1
View 2
Visibility | Enabled
Type | 3 values and horizontal bargraph | View 1, 2nd value | View 2, 2nd value
1st Value | Totalizer 2
2nd Value | Totalizer 2
3rd Value | Totalizer 3
|  | View 1, 3rd value | View 2, 3rd value
Note on B16: Distance from sensor face to zero flow point.
Note on B17: Upper calibration = (Lower calibration point) - (Maximum head)
Note on B32: Totalizer conversion factor = How many custom units per cubic meter. Example: 264.2 gallons per cubic meter.

## LT500 Pump Control
Siemens LT500 Pump Control
| Location: | [7 empty] | Length units
| Date: | 2026-07-29 | [7 empty] | Display, View 1 | Display, View 2
Go to: Quick Start/Quick commissioning/Measurement point/Level
| Measurement Point 1 |  | Level
| Units |  | Inches | [5 empty] | Sensor type
| Sensor type |  | 4-20 mA
Lower calibration point
| Upper calibration point | [8 empty] | View 1, 1st value | View 2, 1st value
| Response rate |  | Slow
| Material type |  | Liquid | [5 empty] | Response rate
Pump control
Go to: Quick start/Pump control/Measurement point 1
| Number of pumps |  | 3 | [6 empty] | View 1, 2nd value | View 2, 2nd value
| Pump control mode |  | Alternate duty assist | [5 empty] | Number of pumps
| Pump number |  | Pump 1 | Pump 2 | Pump 3
| Relay number |  | 1 | 2 | 3
On setpoint
Off setpoint
[10 empty] | View 1, 3rd value | View 2, 3rd Value
| Alarm control | [7 empty] | Pump control mode
Go to: Quick start/Basic control
| Control source |  | Level (point 1)
| Relay output |  |  |  |  | Relay 4 | Relay 5 | Relay 6
On setpoint
| Off setpoint | [8 empty] | View 1, 4th value | View 2, 4th Value
4-20 ma outputs
Go to: Setup/Inputs and outputs
| Current output |  | 1 | 2
| Loop current scale |  | 4-20 mA US | 4-20 mA US
| Process value |  | Level (point 1) | [6 empty] | View 1, 5th value | View 2, 5th Value
Upper range value
Lower range value
Damping value (seconds)
Display setup
| Go to: Setup/Display | [8 empty] | View 1, 6th value | View 2, 6th Value
| View |  | 1 |  |  | 2
| Type |  | 1 value, icons and vertical bargraph |  |  | 6 values
| 1st Value |  | Level (point 1) |  |  | Level (point 2)
| 2nd Value |  | None |  |  | Distance (point 1)
| 3rd Value |  | None |  |  | Current output 1
| 4th Value |  | None |  |  | Level (point 2)
| 5th Value |  | None |  |  | Distance (point 2)
| 6th Value |  | None |  |  | Current output 2
Note on B18: Alternate Duty Assist: Alternates ON and OFF setpoints, and allows multiple pumps to run. MOST COMMON FUNCTION Alternate Duty Backup: alternates ON and OFF setpoints, and allows only one pump to run at a time. Service Ratio Duty Assist: Pumps start based on ratios assigned in P122(1-6). Multiple pumps can run simultaneously. Service Ratio Duty Backup: Pumps start based on ratios assigned in P122(1-6). Only one pump to run at a time. Fixed Duty Assist: Fixed ON and OFF setpoints (no alternation). Allows multiple pumps to run simultaneously. Fixed Duty Backup: Fixed ON and OFF setpoints, Allows only one pump to run at a time. Ex: pump 1 turns off when pump 2 starts.
Note on B27: Cannot be same relay as one used for pump control
Note on B36: Sets process value that corresponds to 20 mA output
Note on B37: Stes process value that corresponds to 0/4 mA out.
Note on B43: Current output 1 and Current output 2 can only be displayed in '6 Value' view.

## LT500 Differential
Siemens LT500 Differential Level
Location: |  |  |  |  | Response rate | Relay 1
Date: | 2026-07-29 | [5 empty] | Display, View 1
Go to: Quick Start/Quick Commissioning
Measurement Point |  | 1 | 2 |  | Length units
Application |  | Level | Level |  |  | Relay 2
Length units |  | Inches | Inches
Sensor Type |  | 4-20 mA | 4-20 mA |  |  |  | View 1, 1st value
Lower calibration point
Upper calibration point
Response rate |  | Medium | Medium |  | Sensor type | Relay 3
Material type |  | Liquid | Liquid
View 1, 2nd value
Relays
Go to: Quick Start/Basic control
Relay |  | Parameter | On setpoint | Off Setpoint | Current Loop Scale | Relay 4
1 |  | Level (point 2)
2 |  | Level difference |  |  |  |  | View 1, 3rd value
3 |  | Level (point 1)
4 |  | Level difference
5 |  | Level (point 1) |  |  | MA out 1 assignment | Relay 5
6 |  | Level (point 2)
4-20 mA outputs | [6 empty] | Display, View 2
Go to: Setup/Inputs and outputs
Current output |  | 1 | 2 |  |  | Relay 6
Current loop scale |  | 4-20 mA US | 4-20 mA US |  | MA out 2 assignment
Process value |  | Level (point 1) | Level Difference
Upper range value | [6 empty] | View 2, 1st value
Lower range value
Damping value (seconds)
Displays
Go to: Setup/Display | [6 empty] | View 2, 2nd value
View |  | 1 | 2
Type |  | 3 values and horizontal bargraph | 3 values and horizontal bargraph
1st Value |  | Level (point 1) | Level difference
2nd Value |  | Level (point 1) | Level (point 1) |  |  |  | View 2, 3rd value
3rd Value |  | Level (point 2) | Level (point 2)
Note on B31: Sets process value that corresponds to 20 mA output
Note on B32: Stes process value that corresponds to 0/4 mA out.
Note on B38: Current output 1 and Current output 2 can only be displayed in '6 value' view.

## LT500 Travelling screen data (hidden)
Response rate
| Sensor 1 | Sensor 2 | [7 empty] | 2
| Slow | Slow | 2
| Medium | Medium
| Fast | Fast
Sensor type
| LR110 | LR110 | 4
| LR120 | LR120
| Probe LU240 | Probe LU240
| 4-20 mA | Generic 4-20 mA
Units
| Meters | Meters | 5
| Centimeters | Centimeters
| Millimeters | Millimeters
| Feet | Feet
| Inches | Inches
| milliamp output scale |  | 4
| 4-20 mA Namur |  | 2
4-20 mA US
4-20 mA( fault=0 mA)
0-20 mA
4-20 ma assignment
| Level (point 1) | Level (point 1) | 1
| Level (point 2) | Level (point 2) | 3
| Level Difference | Level Difference
| Relay mode | 1 |  |  | Process alarms |  |  | Alarm condition
1 | Level difference | 3 |  | 1 | Level (pt 1) | 3 | 1 | Above alarm | 2
| Level (point 1) |  |  |  | Level (pt 2) |  |  | Above warning
| Level (point 2) | 2 |  |  | Level (Diff) |  |  | Below warning
2 | Level difference | 1 |  | 2 | Level (pt 1) | 2 |  | Below alarm
| Level (point 1) | 2 |  |  | Level (pt 2) |  | 2 | Above alarm | 3
| Level (point 2) |  |  |  | Level (Diff) |  |  | Above warning
3 | Level difference | 2 |  | 3 | Level (pt 1) | 2 |  | Below warning
| Level (point 1) |  |  |  | Level (pt 2) |  |  | Below alarm
| Level (point 2) |  |  |  | Level (Diff) |  | 3 | Above alarm | 2
4 | Level difference | 1 |  | 4 | Level (pt 1) |  |  | Above warning
| Level (point 1) |  |  |  | Level (pt 2) |  |  | Below warning
| Level (point 2) |  |  |  | Level (Diff) |  |  | Below alarm
5 | Level difference | 2 |  | 5 | Level (pt 1) |  | 4 | Above alarm
| Level (point 1) |  |  |  | Level (pt 2) |  |  | Above warning
| Level (point 2) |  |  |  | Level (Diff) |  |  | Below warning
6 | Level difference | 3 |  | 6 | Level (pt 1) |  |  | Below alarm
| Level (point 1) |  |  |  | Level (pt 2) |  | 5 | Above alarm
| Level (point 2) |  |  |  | Level (Diff) |  |  | Above warning
Below warning
[7 empty] | 6 | Below alarm
Above alarm
| View 1 |  |  | View 2 |  |  |  | Above warning
| 1 value | 5 |  | 1 value | 5 |  |  | Below warning
| 1 Value and horizontal bargraph |  |  | 1 Value and horizontal bargraph |  |  |  | Below alarm
| 1 value, icons and vertical bargraph |  |  | 1 value, icons and vertical bargraph
| 1 value and trend chart |  |  | 1 value and trend chart
| 3 values and horizontal bargraph |  |  | 3 values and horizontal bargraph
| 3 values, icons and vertical bargraph |  |  | 3 values, icons and vertical bargraph
| 6 values |  |  | 6 values
| 1st value |  |  | 1st value
| Level (point 1) | 1 |  | Level (point 1) | 3
| Level (point 2) |  |  | Level (point 2)
| Level difference |  |  | Level difference
| current output 1 |  |  | current output 1
| current output 2 |  |  | current output 2
| 2nd value |  |  | 2nd value
| Level (point 1) | 1 |  | Level (point 1) | 1
| Level (point 2) |  |  | Level (point 2)
| Level difference |  |  | Level difference
| current output 1 |  |  | current output 1
| current output 2 |  |  | current output 2
| none |  |  | none
| 3rd value | 1 |  | 3rd value
| Level (point 1) | 2 |  | Level (point 1) | 2
| Level (point 2) |  |  | Level (point 2)
| Level difference |  |  | Level difference
| current output 1 |  |  | current output 1
| current output 2 |  |  | current output 2
| none |  |  | none

## HydroRanger 200 OCF
Siemens HydroRanger-200, Multiranger 200, Enviroranger Open Channel Flowmeter
Sensor Type
XPS-10
Date: |  | 2019-09-19
Location: |  | Montpelier WTP
Primary Device: |  | V-notch, 22.5 degree
Max flow (GPM): |  | 100 | Primary Device
V-notch, 22.5 degree
Parameter | Value | Description
P001 | 6 | Operation (Open channel flow)
P002 | 1 | Material: Liquid
P003 | 2 | Rate of change: Medium
P004 | 102 | Sensor type
P005 | 5 | Units of Measure (Inches) | Flow Units
P006 | 74.25 | *Distance from sensor face zero flow level | GPM
P007 | 8.7053 | Head at maximum flow (inches)
P200 | 2 | Output range (4-20 mA)
P201 | 5 | Analog output function (flow rate)
P600 | 1 | Type of primary device | LCD Totalizer Multiplier
P601 | 2.5 | Exponent value | 1
P602 | 0 | Flume width (Palmer Bowlus only) | If flow units GPM, multiply displayed gallons
P603 | 8.7053 | Head at maximum flow (inches) | by this number
P604 | 100 | Maximum flow rate (GPM)
P606 | 2 | Time units | If rate is MGD, .001 makes totalizer read
P608 | 0 | Flow rate units (defined by P604) | Gallons X 1000, (or million gallons X.001)
P620 | 0.6 | Low Flow Cutoff
P630 | 0 | LCD Totalizer multiplier (Reading x multiplier)
P633 | 2 | LCD Totalizer decimal position
P732 | 1.5 | Display Delay (Seconds)
P737 | 1 | Primary display.
P801 | 24 | * Range extension (inches) | Totalizer Decimal Position
P111 | 0 | Relay 1 used for remote totalizer | 2 digits after decimal
P640 | 2 | Remote totalizer multiplier
* Measured variables
Primary Display
Flow rate
Remote Totalizer
Remote totalizer off
Remote Totalizer Multiplier
100
Rate of Change
Rate of change: Medium
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421
Note on D27: Level (in inches) below which totalizer activity is stopped
Note on D28: For an LCD total display in 1,000s of volume units, enter '3' in parameter. Example: Multiplier = 100 Totalizer display = 35 Actual total = 3500
Note on D30: Time (in seconds) between display changes when in toggle mode.
Note on D32: The distance below zero flow condition (PO06) that the level can fall to. This allows water level to fall below 'zero flow' level, without reporting Loss of Echo (LOE).
Note on D34: Example: For a remote totalizer update by 1,000s of volume units, enter '3' in this parameter

## Hydroranger tables (hidden)
Size | Exponent | Head (ft) |  |  |  | Sensor type
1" Parshall flume | 1.55 | 0.76425 | 1 |  | 30 | STH | 100
2" Parshall flume | 1.55 | 0.48868 | 1 |  |  | XPS-10 | 102 | 2
3" Parshall Flume | 1.557 | 0.38323 | 1 |  |  | XPS-15 | 104
6" Parshall Flume | 1.58 | 0.24472 | 1 |  |  | XRS-5 | 112
9" Parshall flume | 1.53 | 0.18005 | 1 |  |  | Milliamp input | 250
12" Parshall flume | 1.522 | 0.14998 | 1
18" Parshall flume | 1.538 | 0.11751 | 1 |  |  | Flow units | Multiplier | Time
24" Parshall flume | 1.55 | 0.099246 | 1 |  |  | GPM | 1 | 2 | Maximum flow rate (GPM) | 1
30" Parshall flume | 1.559 | 0.087159 | 1 |  |  | MGD | 0.00144 | 4 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 0.078427 | 1
4'Parshall flume | 1.578 | 0.066636 | 1
5' Parshall flume | 1.587 | 0.058793 | 1 |  |  | Totalizer scale factor
6' Parshall flume | 1.595 | 0.053197 | 1 |  |  | 0.001 | -3 | 4
8' Parshall flume | 1.607 | 0.045462 | 1 |  |  | 0.01 | -2
10' Parshall flume | 1.6 | 0.039398 | 1 |  |  | 0.1 | -1
12' Parshall flume | 1.55 | 0.031773 | 1 |  |  | 1 | 0
4" Leopold Lagco flume | 1.547 | 0.42256 | 1 |  |  | 10 | 1
6" Leopold Lagco flume | 1.547 | 0.32917 | 1 |  |  | 100 | 2
8" Leopold Lagco flume | 1.547 | 0.2757 | 1 |  |  | 1000 | 3
10" Leopold Lagco flume | 1.547 | 0.24031 | 1 |  |  | 10000 | 4
12" Leopold Lagco flume | 1.547 | 0.21477 | 1 |  |  | 100000 | 5
15" Leopold Lagco flume | 1.547 | 0.18779 | 1 |  |  | 1000000 | 6
18" Leopold Lagco flume | 1.547 | 0.16733 | 1 |  |  | 10000000 | 7
21" Leopold Lagco flume | 1.547 | 0.15214 | 1
24" Leopold Lagco flume | 1.547 | 0.14012 | 1 |  |  | Totalizer Decimal Position
Trapezoidal, small 60 deg | 2.58 | 0.47152 | 1 |  |  | No decimal | 0 | 3
Trapezoidal, large 60 deg V | 2.58 | 0.47152 | 1 |  |  | 1 digit after decimal | 1
Trapezoidal, XL 60 degree V | 2.6363 | 0.47176 | 1 |  |  | 2 digits after decimal | 2
Trapezoidal, 3' 60 degree V | 2.5804 | 0.50114 | 1 |  |  | 3 digits after decimal | 3
V-notch, 22.5 degree | 2.5 | 0.72544 | 1
V-notch, 30 degree | 2.5 | 0.6415 | 1
V-notch, 45 degree | 2.5 | 0.54101 | 1 |  |  | Primary Display
V-notch, 60 degree | 2.5 | 0.47367 | 1 |  |  | Flow rate | 1 | 1
V-notch, 90 degree | 2.5 | 0.38019 | 1 |  |  | Flow total | 2
V-notch, 120 degree | 2.5 | 0.30522 | 1 |  |  | Toggle flow/total, 3 seconds each | 3
1' Rectangular weir without end contractions | 1.5 | 0.16478 | 1
1.5' Rectangular weir without end contractions | 1.5 | 0.12577 | 1
2' Rectangular weir without end contractions | 1.5 | 0.10383 | 1 |  |  | Remote totalizer
2.5' Rectangular weir without end contractions | 1.5 | 0.089481 | 1 |  |  | Remote totalizer on | 40 | 2
3' Rectangular weir without end contractions | 1.5 | 0.07923 | 1 |  |  | Remote totalizer off | 0
4' Rectangular weir without end contractions | 1.5 | 0.065408 | 1
5' Rectangular weir without end contractions | 1.5 | 0.056364 | 1
6' Rectangular weir without end contractions | 1.5 | 0.049915 | 1
8' Rectangular weir without end contractions | 1.5 | 0.041195 | 1 |  |  | Remote totalizer scale factor
10' Rectangular weir without end contractions | 1.5 | 0.035501 | 1 |  |  | 0.001 | -3 | 6
4" Palmer Bowlus Flume |  | 0.34507 | 2 | 4 |  | 0.01 | -2
6" Palmer Bowlus Flume |  | 0.30438 | 2 | 6 |  | 0.1 | -1
8" Palmer Bowlus Flume |  | 0.2768 | 2 | 8 |  | 1 | 0
10" Palmer Bowlus Flume |  | 0.25541 | 2 | 10 |  | 10 | 1
12" Palmer Bowlus Flume |  | 0.23787 | 2 | 12 |  | 100 | 2
15" Palmer Bowlus Flume |  | 0.21639 | 2 | 15 |  | 1000 | 3
18" Palmer Bowlus Flume |  | 0.19994 | 2 | 18 |  | 10000 | 4
21" Palmer Bowlus Flume |  | 0.18425 | 2 | 21 |  | 100000 | 5
27" Palmer Bowlus Flume |  | 0.16165 | 2 | 27 |  | 1000000 | 6
30" Palmer Bowlus Flume |  | 0.14258 | 2 | 30 |  | 10000000 | 7
36" Palmer Bowlus Flume |  | 0.12053 | 2 | 36
42" Palmer Bowlus Flume |  | 0.1069 | 2 | 42
48" Palmer Bowlus Flume |  | 0.091912 | 2 | 48 |  | Rate of Change
60" Palmer Bowlus Flume |  | 0.047561 | 2 | 60 |  | Rate of Change: Slow | 1 | 2 | Rate of change: Slow
72" Palmer Bowlus Flume |  | 0.011661 | 2 | 72 |  | Rate of Change: Medium | 2 |  | Rate of change: Medium
[6 empty] | Rate of Change: Fast | 3 |  | Rate of change: Fast

## OCM-III tables (hidden)
Size | Exponent | Head (ft)
1" Parshall flume | 1.55 | 2.6672 | 0 | Head at max flow (inches) | 7
2" Parshall flume | 1.55 | 1.7054 | 0 | Head at max flow (inches)
3" Parshall Flume | 1.557 | 1.3299 | 0 | Head at max flow (inches)
6" Parshall Flume | 1.58 | 0.83401 | 0 | Head at max flow (inches)
9" Parshall flume | 1.53 | 0.63871 | 0 | Head at max flow (inches)
12" Parshall flume | 1.522 | 0.5356 | 0 | Head at max flow (inches)
18" Parshall flume | 1.538 | 0.41411 | 0 | Head at max flow (inches) |  | Flow units | Multiplier | Time
24" Parshall flume | 1.55 | 0.34636 | 0 | Head at max flow (inches) |  | GPM | 1 | 3 | Maximum flow rate (GPM) | 1
30" Parshall flume | 1.559 | 0.30199 | 0 | Head at max flow (inches) |  | MGD | 0.00144 | 5 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 0.27023 | 0 | Head at max flow (inches)
4'Parshall flume | 1.578 | 0.22745 | 0 | Head at max flow (inches)
5' Parshall flume | 1.587 | 0.19929 | 0 | Head at max flow (inches) |  | Totalizer scale factor
6' Parshall flume | 1.595 | 0.17922 | 0 | Head at max flow (inches) |  | 0.001 | 0 | 6
8' Parshall flume | 1.607 | 0.15178 | 0 | Head at max flow (inches) |  | 0.01 | 1
10' Parshall flume | 1.6 | 0.13223 | 0 | Head at max flow (inches) |  | 0.1 | 2
12' Parshall flume | 1.55 | 0.11089 | 0 | Head at max flow (inches) |  | 1 | 3
4" Leopold Lagco flume | 1.547 | 1.4783 | 0 | Head at max flow (inches) |  | 10 | 4
6" Leopold Lagco flume | 1.547 | 1.1515 | 0 | Head at max flow (inches) |  | 100 | 5
8" Leopold Lagco flume | 1.547 | 0.96451 | 0 | Head at max flow (inches) |  | 1000 | 6
10" Leopold Lagco flume | 1.547 | 0.8407 | 0 | Head at max flow (inches)
12" Leopold Lagco flume | 1.547 | 0.75136 | 0 | Head at max flow (inches)
15" Leopold Lagco flume | 1.547 | 0.65537 | 0 | Head at max flow (inches)
18" Leopold Lagco flume | 1.547 | 0.58539 | 0 | Head at max flow (inches)
21" Leopold Lagco flume | 1.547 | 0.53225 | 0 | Head at max flow (inches)
24" Leopold Lagco flume | 1.547 | 0.4902 | 0 | Head at max flow (inches) |  | Flow rate Decimal Position
Trapezoidal, small 60 deg | 2.58 | 0.99911 | 0 | Head at max flow (inches) |  | No decimal | 0 | 3
Trapezoidal, large 60 deg V | 2.58 | 0.99911 | 0 | Head at max flow (inches) |  | 1 digit after decimal | 1
Trapezoidal, XL 60 degree V | 2.6363 | 0.99961 | 0 | Head at max flow (inches) |  | 2 digits after decimal | 2
Trapezoidal, 3' 60 degree V | 2.5804 | 1.0617 | 0 | Head at max flow (inches) |  | 3 digits after decimal | 3
V-notch, 22.5 degree | 2.5 | 1.5745 | 0 | Head at max flow (inches) |  | 4 digits after decimal | 4
V-notch, 30 degree | 2.5 | 1.3923 | 0 | Head at max flow (inches)
V-notch, 45 degree | 2.5 | 1.1742 | 0 | Head at max flow (inches)
V-notch, 60 degree | 2.5 | 1.0281 | 0 | Head at max flow (inches)
V-notch, 90 degree | 2.5 | 0.82518 | 0 | Head at max flow (inches)
V-notch, 120 degree | 2.5 | 0.66245 | 0 | Head at max flow (inches)
1' Rectangular weir without end contractions | 1.5 | 0.59953 | 0 | Head at max flow (inches)
1.5' Rectangular weir without end contractions | 1.5 | 0.4576 | 0 | Head at max flow (inches)
2' Rectangular weir without end contractions | 1.5 | 0.37777 | 0 | Head at max flow (inches)
2.5' Rectangular weir without end contractions | 1.5 | 0.32556 | 0 | Head at max flow (inches) |  | Temperature units
3' Rectangular weir without end contractions | 1.5 | 0.28827 | 0 | Head at max flow (inches) |  | Deg C | 0 | 2
4' Rectangular weir without end contractions | 1.5 | 0.23798 | 0 | Head at max flow (inches) |  | Deg F | 1
5' Rectangular weir without end contractions | 1.5 | 0.20507 | 0 | Head at max flow (inches)
6' Rectangular weir without end contractions | 1.5 | 0.18161 | 0 | Head at max flow (inches)
8' Rectangular weir without end contractions | 1.5 | 0.14988 | 0 | Head at max flow (inches) |  | Remote relay
10' Rectangular weir without end contractions | 1.5 | 0.12917 | 0 | Head at max flow (inches) |  | Yes | 35 | 2
4" Palmer Bowlus Flume | 4 | 0.93841 | 10 | Head at max flow (inches) |  | no | 0
6" Palmer Bowlus Flume | 6 | 0.85226 | 10 | Head at max flow (inches)
8" Palmer Bowlus Flume | 8 | 0.76346 | 10 | Head at max flow (inches)
10" Palmer Bowlus Flume | 10 | 0.71603 | 10 | Head at max flow (inches)
12" Palmer Bowlus Flume | 12 | 0.67619 | 10 | Head at max flow (inches)
15" Palmer Bowlus Flume | 15 | 0.62754 | 10 | Head at max flow (inches)
18" Palmer Bowlus Flume | 18 | 0.58825 | 10 | Head at max flow (inches)
21" Palmer Bowlus Flume | 21 | 0.56029 | 10 | Head at max flow (inches)
27" Palmer Bowlus Flume | 27 | 0.50765 | 10 | Head at max flow (inches)
30" Palmer Bowlus Flume | 30 | 0.48681 | 10 | Head at max flow (inches)
36" Palmer Bowlus Flume | 36 | 0.4506 | 10 | Head at max flow (inches)
42" Palmer Bowlus Flume | 42 | 0.41812 | 10 | Head at max flow (inches)
48" Palmer Bowlus Flume | 48 | 0.3907 | 10 | Head at max flow (inches)
60" Palmer Bowlus Flume | 60 | 0.33497 | 10 | Head at max flow (inches)
72" Palmer Bowlus Flume | 72 | 0.28682 | 10 | Head at max flow (inches)
1' Rectangular weir, contracted | 12 |  | 8 | Head at max flow (inches), See tables
1.5' Rectangular weir, contracted | 18 |  | 8 | Head at max flow (inches), See tables
2' Rectangular weir, contracted | 24 |  | 8 | Head at max flow (inches), See tables
2.5' Rectangular weir, contracted | 30 |  | 8 | Head at max flow (inches), See tables
3' Rectangular weir, contracted | 36 |  | 8 | Head at max flow (inches), See tables
4' Rectangular weir, contracted | 48 |  | 8 | Head at max flow (inches), See tables
5' Rectangular weir, contracted | 60 |  | 8 | Head at max flow (inches), See tables
6' Rectangular weir, contracted | 72 |  | 8 | Head at max flow (inches), See tables
8' Rectangular weir, contracted | 96 |  | 8 | Head at max flow (inches), See tables
10' Rectangular weir, contracted | 120 |  | 8 | Head at max flow (inches), See tables

## Hydro-HMI tables (hidden)
Size | Exponent | Head (ft) |  |  |  |  | Sensor type
1" Parshall flume | 1.55 | 2.1587 | Exponential |  | 4 |  | STH | 100
2" Parshall flume | 1.55 | 1.3803 | Exponential |  |  |  | XPS-10 | 102 | 3
3" Parshall Flume | 1.557 | 1.0774 | Exponential |  |  |  | XPS-15 | 104
6" Parshall Flume | 1.58 | 0.67772 | Exponential |  |  |  | XRS-5 | 112
9" Parshall flume | 1.53 | 0.51551 | Exponential |  |  |  | Milliamp input | 250
12" Parshall flume | 1.522 | 0.4318 | Exponential
18" Parshall flume | 1.538 | 0.33461 | Exponential |  |  |  | Flow units | Multiplier | Time
24" Parshall flume | 1.55 | 0.28032 | Exponential |  |  | Minutes | GPM | 1 | 2 | Maximum flow rate (GPM) | 1
30" Parshall flume | 1.559 | 0.24471 | Exponential |  |  | Days | MGD | 0.00144 | 4 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 0.21918 | Exponential
4'Parshall flume | 1.578 | 0.18478 | Exponential
5' Parshall flume | 1.587 | 0.16209 | Exponential |  |  |  | Totalizer scale factor
6' Parshall flume | 1.595 | 0.14592 | Exponential |  |  |  | 0.001 | -3 | 4
8' Parshall flume | 1.607 | 0.12377 | Exponential |  |  |  | 0.01 | -2
10' Parshall flume | 1.6 | 0.10773 | Exponential |  |  |  | 0.1 | -1
12' Parshall flume | 1.55 | 0.089745 | Exponential |  |  |  | 1 | 0
4" Leopold Lagco flume | 1.547 | 1.1959 | Exponential |  |  |  | 10 | 1
6" Leopold Lagco flume | 1.547 | 0.93162 | Exponential |  |  |  | 100 | 2
8" Leopold Lagco flume | 1.547 | 0.78031 | Exponential |  |  |  | 1000 | 3
10" Leopold Lagco flume | 1.547 | 0.68014 | Exponential |  |  |  | 10000 | 4
12" Leopold Lagco flume | 1.547 | 0.60786 | Exponential |  |  |  | 100000 | 5
15" Leopold Lagco flume | 1.547 | 0.53042 | Exponential |  |  |  | 1000000 | 6
18" Leopold Lagco flume | 1.547 | 0.47359 | Exponential |  |  |  | 10000000 | 7
21" Leopold Lagco flume | 1.547 | 0.4306 | Exponential
24" Leopold Lagco flume | 1.547 | 0.39658 | Exponential |  |  |  | Totalizer Decimal Position
Trapezoidal, small 60 deg | 2.58 | 0.87988 | Exponential |  |  |  | No decimal | 0 | 1
Trapezoidal, large 60 deg V | 2.58 | 0.87988 | Exponential |  |  |  | 1 digit after decimal | 1
Trapezoidal, XL 60 degree V | 2.6363 | 0.88032 | Exponential |  |  |  | 2 digits after decimal | 2
Trapezoidal, 3' 60 degree V | 2.5804 | 0.93505 | Exponential |  |  |  | 3 digits after decimal | 3
V-notch, 22.5 degree | 2.5 | 1.381 | Exponential
V-notch, 30 degree | 2.5 | 1.2212 | Exponential
V-notch, 45 degree | 2.5 | 1.0299 | Exponential |  |  |  | Primary Display
V-notch, 60 degree | 2.5 | 0.90171 | Exponential |  |  |  | Flow rate | 1 | 2
V-notch, 90 degree | 2.5 | 0.72375 | Exponential |  |  |  | Flow total | 2
V-notch, 120 degree | 2.5 | 0.58103 | Exponential |  |  |  | Toggle flow/total, 3 seconds each | 3
1' Rectangular weir without end contractions | 1.5 | 0.48182 | Exponential
1.5' Rectangular weir without end contractions | 1.5 | 0.36775 | Exponential
2' Rectangular weir without end contractions | 1.5 | 0.3036 | Exponential |  |  |  | Remote totalizer
2.5' Rectangular weir without end contractions | 1.5 | 0.26164 | Exponential |  |  |  | Remote totalizer on | 40 | 2
3' Rectangular weir without end contractions | 1.5 | 0.23167 | Exponential |  |  |  | Remote totalizer off | 0
4' Rectangular weir without end contractions | 1.5 | 0.19125 | Exponential
5' Rectangular weir without end contractions | 1.5 | 0.16481 | Exponential
6' Rectangular weir without end contractions | 1.5 | 0.14595 | Exponential
8' Rectangular weir without end contractions | 1.5 | 0.12046 | Exponential |  |  |  | Remote totalizer scale factor
10' Rectangular weir without end contractions | 1.5 | 0.10381 | Exponential |  |  |  | 0.001 | -3 | 1
4" Palmer Bowlus Flume |  | 0.79277 | Palmer Bowlus | 4 |  |  | 0.01 | -2
6" Palmer Bowlus Flume |  | 0.71649 | Palmer Bowlus | 6 |  |  | 0.1 | -1
8" Palmer Bowlus Flume |  | 0.64431 | Palmer Bowlus | 8 |  |  | 1 | 0
10" Palmer Bowlus Flume |  | 0.60289 | Palmer Bowlus | 10 |  |  | 10 | 1
12" Palmer Bowlus Flume |  | 0.56821 | Palmer Bowlus | 12 |  |  | 100 | 2
15" Palmer Bowlus Flume |  | 0.52571 | Palmer Bowlus | 15 |  |  | 1000 | 3
18" Palmer Bowlus Flume |  | 0.49194 | Palmer Bowlus | 18 |  |  | 10000 | 4
21" Palmer Bowlus Flume |  | 0.46806 | Palmer Bowlus | 21 |  |  | 100000 | 5
27" Palmer Bowlus Flume |  | 0.42227 | Palmer Bowlus | 27 |  |  | 1000000 | 6
30" Palmer Bowlus Flume |  | 0.40245 | Palmer Bowlus | 30 |  |  | 10000000 | 7
36" Palmer Bowlus Flume |  | 0.37004 | Palmer Bowlus | 36
42" Palmer Bowlus Flume |  | 0.3418 | Palmer Bowlus | 42
48" Palmer Bowlus Flume |  | 0.31737 | Palmer Bowlus | 48 |  |  | Rate of Change
60" Palmer Bowlus Flume |  | 0.26473 | Palmer Bowlus | 60 |  |  | Slow | Slow | 2
72" Palmer Bowlus Flume |  | 0.21962 | Palmer Bowlus | 72 |  |  | Medium | Medium
[7 empty] | Fast | Fast

## Hydro-200-HMI OCF
Siemens HydroRanger-200 HMI Open Channel Flowmeter
Sensor Type
XPS-15
Date:
Location:
Primary Device: | 6" Parshall Flume
Max flow (GPM): | 500
Rate of Change
Go to: Wizards/Quick Start/QS Flow |  | Medium
Description | Value
Measurement Point Selector | 1
Transducer | XPS-15
Temperature Source | Transducer | Primary Device
Units | Inches | 6" Parshall Flume
Empty (sensor face to zero elevation) | 40
Span | 8.1327
Response Rate (slow, med, fast) | Medium
Primary Measuring Device | Exponential
Flow Exponent | 1.58
Max Head | 8.1327 | Flow Rate Units
Zero Head | 0 | GPM
Max Flow | 500
Flow Rate Decimal (0,1,2,3) | 2
Flow Time Units (Sec ,Min, Hour, Day) | Minutes
Low Flow Cutoff (inches) | 0.2 | LCD Totalizer Multiplier
1
Go to: Setup/Totalizers
Totalizer Decimal (0,1,2,3) | 0 | If flow units GPM, multiply displayed gallons
Totalizer Multiplier | 1 | by this number
To assign relay for remote totalizer, Go to: |  | If rate is MGD, .001 makes totalizer read
Setup/Relays/Basic Setup/Relay Function | Totalizer | Gallons X 1000, (or million gallons X.001)
Setup/Other Control/External Totalizer/Multiplier | 0.001
Cycle power for all parameters to take effect
Remote Totalizer Multiplier
0.001
Enter the factor by which the volume is
divided, prior to remote totalizer actuation
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421

## LUT400 tables (hidden)
Size | Exponent | Head (ft) |  |  |  | Sensor type
1" Parshall flume | 1.55 | 3.376 | Exponential |  | 1 | STH | 100
2" Parshall flume | 1.55 | 2.1587 | Exponential |  |  | XPS-10 | 102 | 3
3" Parshall Flume | 1.557 | 1.6816 | Exponential |  |  | XPS-15 | 104
6" Parshall Flume | 1.58 | 1.0509 | Exponential |  |  | XRS-5 | 112
9" Parshall flume | 1.53 | 0.81094 | Exponential |  |  | Milliamp input | 250
12" Parshall flume | 1.522 | 0.68088 | Exponential
18" Parshall flume | 1.538 | 0.52512 | Exponential |  | LUT-400 | Flow units | Multiplier | Time
24" Parshall flume | 1.55 | 0.4384 | Exponential |  |  | GPM (GAL/MIN) | 1 | 2 | Maximum flow at 20 mA (GPM) | 2
30" Parshall flume | 1.559 | 0.38172 | Exponential |  |  | MGD (MMGAL/D) | 0.00144 | 4 | Maximum flow at 20 mA (MGD)
3' Parshall flume | 1.566 | 0.34122 | Exponential
4' Parshall flume | 1.578 | 0.2867 | Exponential
5' Parshall flume | 1.587 | 0.25087 | Exponential |  |  | Totalizer scale factor
6' Parshall flume | 1.595 | 0.22534 | Exponential |  |  | 0.001 | -3 | 5
8' Parshall flume | 1.607 | 0.19051 | Exponential |  |  | 0.01 | -2
10' Parshall flume | 1.6 | 0.16614 | Exponential |  |  | 0.1 | -1
12' Parshall flume | 1.55 | 0.14035 | Exponential |  |  | 1 | 0
4" Leopold Lagco flume | 1.547 | 1.872 | Exponential |  |  | 10 | 1
6" Leopold Lagco flume | 1.547 | 1.4582 | Exponential |  |  | 100 | 2
8" Leopold Lagco flume | 1.547 | 1.2214 | Exponential |  |  | 1000 | 3
10" Leopold Lagco flume | 1.547 | 1.0646 | Exponential |  |  | 10000 | 4
12" Leopold Lagco flume | 1.547 | 0.95147 | Exponential |  |  | 100000 | 5
15" Leopold Lagco flume | 1.547 | 0.82953 | Exponential |  |  | 1000000 | 6
18" Leopold Lagco flume | 1.547 | 0.74129 | Exponential |  |  | 10000000 | 7
21" Leopold Lagco flume | 1.547 | 0.67401 | Exponential
24" Leopold Lagco flume | 1.547 | 0.62075 | Exponential |  |  | Totalizer Decimal Position
Trapezoidal, small 60 deg | 2.58 | 1.1511 | Exponential |  |  | No decimal | 0 | 1
Trapezoidal, large 60 deg V | 2.58 | 1.1511 | Exponential |  |  | 1 digit after decimal | 1
Trapezoidal, XL 60 degree V | 2.6363 | 1.1516 | Exponential |  |  | 2 digits after decimal | 2
Trapezoidal, 3' 60 degree V | 2.5804 | 1.2232 | Exponential |  |  | 3 digits after decimal | 3
V-notch, 22.5 degree | 2.5 | 1.8222 | Exponential
V-notch, 30 degree | 2.5 | 1.6114 | Exponential
V-notch, 45 degree | 2.5 | 1.359 | Exponential |  |  | Primary Display
V-notch, 60 degree | 2.5 | 1.1898 | Exponential |  |  | Flow rate | 1 | 2
V-notch, 90 degree | 2.5 | 0.955 | Exponential |  |  | Flow total | 2
V-notch, 120 degree | 2.5 | 0.76667 | Exponential |  |  | Toggle flow/total, 3 seconds each | 3
1' Rectangular weir without end contractions | 1.5 | 0.76484 | Exponential
1.5' Rectangular weir without end contractions | 1.5 | 0.58377 | Exponential
2' Rectangular weir without end contractions | 1.5 | 0.48193 | Exponential |  |  | Remote totalizer
2.5' Rectangular weir without end contractions | 1.5 | 0.41533 | Exponential |  |  | Remote totalizer on | 40 | 2
3' Rectangular weir without end contractions | 1.5 | 0.36775 | Exponential |  |  | Remote totalizer off | 0
4' Rectangular weir without end contractions | 1.5 | 0.3036 | Exponential
5' Rectangular weir without end contractions | 1.5 | 0.26162 | Exponential
6' Rectangular weir without end contractions | 1.5 | 0.23169 | Exponential
8' Rectangular weir without end contractions | 1.5 | 0.19121 | Exponential |  |  | Remote totalizer scale factor
10' Rectangular weir without end contractions | 1.5 | 0.16478 | Exponential |  |  | 0.001 | -3 | 6
4" Palmer Bowlus Flume |  | 1.1322 | Palmer Bowlus | 4 |  | 0.01 | -2
6" Palmer Bowlus Flume |  | 1.0338 | Palmer Bowlus | 6 |  | 0.1 | -1
8" Palmer Bowlus Flume |  | 0.92175 | Palmer Bowlus | 8 |  | 1 | 0
10" Palmer Bowlus Flume |  | 0.86661 | Palmer Bowlus | 10 |  | 10 | 1
12" Palmer Bowlus Flume |  | 0.82012 | Palmer Bowlus | 12 |  | 100 | 2
15" Palmer Bowlus Flume |  | 0.76364 | Palmer Bowlus | 15 |  | 1000 | 3
18" Palmer Bowlus Flume |  | 0.71709 | Palmer Bowlus | 18 |  | 10000 | 4
21" Palmer Bowlus Flume |  | 0.68294 | Palmer Bowlus | 21 |  | 100000 | 5
27" Palmer Bowlus Flume |  | 0.62154 | Palmer Bowlus | 27 |  | 1000000 | 6
30" Palmer Bowlus Flume |  | 0.59895 | Palmer Bowlus | 30 |  | 10000000 | 7
36" Palmer Bowlus Flume |  | 0.55745 | Palmer Bowlus | 36
42" Palmer Bowlus Flume |  | 0.5196 | Palmer Bowlus | 42
48" Palmer Bowlus Flume |  | 0.48823 | Palmer Bowlus | 48 |  | Rate of Change
60" Palmer Bowlus Flume |  | 0.42819 | Palmer Bowlus | 60 |  | Slow | Slow | 2
72" Palmer Bowlus Flume |  | 0.37598 | Palmer Bowlus | 72 |  | Medium | Medium
[6 empty] | Fast | Fast

## H200 Rect. weir end contraction
Siemens HydroRanger-200, Multiranger 200, Enviroranger Open Channel Flowmeter
Sensor Type
XRS-5
Date:
Location:
Primary Device: |  | 1' Rectangular weir with end contractions
Max flow (MGD): |  | 1 | Primary Device
1' Rectangular weir with end contractions
Parameter | Value | Description
P001 | 6 | Operation (Open channel flow)
P002 | 1 | Material: Liquid
P003 | 1 | Rate of Change: Slow
P004 | 112 | Sensor type
P005 | 5 | Units of Measure (Inches) | Rate of Change
P006 |  | *Distance from sensor face zero flow level | Rate of Change: Slow
P007 | 12.0024 | Head at maximum flow (inches)
P200 | 2 | Output range (4-20 mA)
P201 | 5 | Analog output function (flow rate)
P600 | 4 | Type of primary device | LCD Totalizer Multiplier
P601 | NA | Exponent value | 100
P602 | NA | Flume width (Palmer Bowlus only)
P603 | 12.0024 | Head at maximum flow (inches)
P604 | 1 | Maximum flow rate (MGD
P606 | 4 | Time units | Rate is MGD, .001 makes totalizer read
P608 | 0 | Flow rate units (defined by P604) | Gallons X 1000, (or million gallons X.001)
P620 | 0.2 | Low Flow Cutoff
P630 | 2 | LCD Totalizer multiplier (Reading x multiplier)
P633 | 2 | LCD Totalizer decimal position
P732 | 5 | Display Delay (Seconds)
P737 | 1 | Primary display.
P801 |  | * Range extension (inches) | Totalizer Decimal Position
P111 | 40 | Relay 1 used for remote totalizer | 2 digits after decimal
P640 | -1 | Remote totalizer multiplier
* Measured variables
Primary Display
Flow rate
Remote Totalizer
Remote totalizer on
Remote Totalizer Multiplier
0.1
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Level/Flow Conversion Table
Max level (in) | 12.0024
Max flow mgd | 1
Weir crest (ft) | 1
Index# | P610 (in) | P611 (MGD)
1 | 0 | 0
2 | 0.77435 | 0.03482
3 | 1.1615 | 0.063551
4 | 1.5487 | 0.097199
5 | 1.9359 | 0.13494
6 | 2.323 | 0.1762
7 | 2.7102 | 0.22055
8 | 3.0974 | 0.26764
9 | 3.4846 | 0.31718
10 | 3.8717 | 0.36894
11 | 4.2589 | 0.42271
12 | 4.6461 | 0.4783
13 | 5.0333 | 0.53554
14 | 5.4204 | 0.59429
15 | 5.8076 | 0.65442
16 | 6.1948 | 0.71579
17 | 6.582 | 0.77829
18 | 6.9691 | 0.84181
19 | 7.3563 | 0.90627
20 | 7.7435 | 0.97155
21 | 8.1307 | 1.0376
22 | 8.5178 | 1.1043
23 | 8.905 | 1.1715
24 | 9.2922 | 1.2393
25 | 9.6794 | 1.3075
26 | 10.0665 | 1.376
27 | 10.4537 | 1.4449
28 | 10.8409 | 1.514
29 | 11.228 | 1.5832
30 | 11.6152 | 1.6526
31 | 12.0024 | 1.722
Instructions
1. Select sensor type, weir size and multiplier from lists.
2. Insert maximum flow rate in cell D8 (in MGD).
3. Hit 'Calculate Max Head' button
4. Program all parameters, including level and flow table into controller.
Note on D27: Level (in inches) below which totalizer activity is stopped
Note on D28: For an LCD total display in 1,000s of volume units, enter '3' in parameter. Example: Multiplier = 100 Totalizer display = 35 Actual total = 3500
Note on D30: Time (in seconds) between display changes when in toggle mode.
Note on D32: The distance below zero flow condition (PO06) that the level can fall to. This allows water level to fall below 'zero flow' level, without reporting Loss of Echo (LOE).
Note on D34: Example: For a remote totalizer update by 1,000s of volume units, enter '3' in this parameter
Calculate Max Head

## H200-HMI Contracted weir
Siemens HydroRanger-200 HMI Open Channel Flowmeter
Instructions
Date: |  | 1. Enter weir width (in feet)
Location: |  | 2. Enter max flow rate (in MGD)
Contracted Rectangular Weir Width (Ft) | 4 | 3. Enter head at max flow (in inches)
Max flow (MGD): | 1.89
Head at Max flow: (inches) | 9
Level vs Flow table
Description | Value | Number | Level (in.) | Flow (MGD)
Setup/Sensor |  | 1 | 0 | 0
Units | Inches | 2 | 0.5625 | 0.087156
Level Selector | Level 1 | 3 | 0.84375 | 0.15993
Sensor Mode | Flow rate in open channel | 4 | 1.125 | 0.24593
Transducer Selector | Transducer 1 | 5 | 1.4062 | 0.3433
Transducer |  | 6 | 1.6875 | 0.45075
Material | Liquid | 7 | 1.9688 | 0.56733
Setup/Calibration |  | 8 | 2.25 | 0.69233
Level Selector | Level 1 | 9 | 2.5312 | 0.82514
Span | 9 | 10 | 2.8125 | 0.96527
Transducer Selector | Transducer 1 | 11 | 3.0938 | 1.1123
Empty |  | 12 | 3.375 | 1.2659
Sensor Offset | 0 | 13 | 3.6562 | 1.4257
Blanking | 11.8 | 14 | 3.9375 | 1.5914
Range Extension |  | 15 | 4.2188 | 1.7628
Setup/Rate |  | 16 | 4.5 | 1.9397
Transducer Selector | 1 | 17 | 4.7812 | 2.1218
Response Rate |  | 18 | 5.0625 | 2.309
Setup/Flow |  | 19 | 5.3438 | 2.501
Transducer Selector | 1 | 20 | 5.625 | 2.6978
Primary Measuring Device | Universal linear flow calc | 21 | 5.9062 | 2.8992
Auto Zero Head | 0 | 22 | 6.1875 | 3.105
Basic Setup |  | 23 | 6.4688 | 3.3151
Flow Exponent | NA | 24 | 6.75 | 3.5294
Maximum Head | 9 | 25 | 7.0312 | 3.7477
Maximum Flow | 1.89 | 26 | 7.3125 | 3.97
Flow Time Units | Day | 27 | 7.5938 | 4.1962
Zero Head | 0 | 28 | 7.875 | 4.4261
Flowrate Decimal |  | 29 | 8.1562 | 4.6596
Flowrate Units | MUSGAL/D | 30 | 8.4375 | 4.8968
Low Flow Cutoff |  | 31 | 8.7188 | 5.1374
PMD Dimension | NA | 32 | 9 | 5.3814
Universal Head vs Flow | Data from table to right
Setup/totalizer
Totalizer Decimal (0,1,2,3)
Totalizer Multiplier
To assign relay for remote totalizer, Go to:
Setup/Relays/Basic Setup/Relay Function | Totalizer
Setup/Other Control/External Totalizer/Multiplier
Note on A15: Select sensor model# (XRS-5, XPS-10, XPS-15, etc)
Note on A21: Distance from transducer face to zero flow point
Note on A22: May be used to fine-tune calibration. Refer to manual
Note on A23: Space belo transducer face that cannot be measured. Default= 11.8"
Note on A24: Distance below zero level where echoes are still valid (no loss of echo fault).
Note on A26: Select transducer 1 or 2. Always select transducer 1 for single channel flowmeter
Note on A27: Fast, medium or slow response rate.
Note on A31: Use this parameter when the reported head is consistently high or low by a fixed amount. The following is the procedure with “head” steady: Step 1. Measure the actual head (with a tape measure or solid rule) Step 2. Enter the actual head value. The deviation between the entered 2.2.4. Empty value and the calibrated Empty value is stored in 2.11.1.9. Offset Correction. Alternatively, the 2.2.4. Empty can be corrected directly.
Note on A34: Level above weir plate that represents maximum flow.
Note on A37: The distance above empty. For rectaugular weir, set to 0".
Note on A40: Eliminates totalizer activity for flows at, or below the cutoff value (in inches)
Note on A45: Enter the factor (powers of 10 only) by which actual volume is divided, prior to display on the LCD. Use a value such that the eight-digit totalizer doesn’t roll over between readings. Example: If flow units are MGD, to display total in thousands of gallons (0.001 Million Gallons), set Multiplier to 0.001.
Note on A48: Enter the factor (powers of 10 only) by which actual volume is divided, prior to Remote Totalizer count increment. Example: If flow units are MGD, to activate relay every 1,000 gallons (0.001 Million Gallons), set multiplier to 0.001.

## LUT-400 OCF
LUT400 Open Channel Flowmeter
Sensor Type
XPS-15
Date:
Location:
Primary Device: | 1" Parshall flume
Max flow (GPM): | 1000
Rate of Change
Go to: Wizards/Quick Start/QS Flow |  | Medium
Description | Value
Transducer | XPS-15
Temperature Source | Transducer
Primary Measuring Device | Exponential | Primary Device
Method of Flow Calculation | Ratiometric | 1" Parshall flume
Units | Inches
High Calibration Point | -0.51144
Low Calibration Point (sensor face to zero elevation) | 40
Response Rate (slow, med, fast) | Medium
Flow Exponent | 1.55
Maximum Head | 40.5114
Zero Head offset | 0 | Flow Rate Units
Flow Rate Units | MGD (MMGAL/D) | MGD (MMGAL/D)
Maximum flow at 20 mA (MGD) | 1.44
Flow Rate Decimal (0,1,2,3) | 0
Low Flow Cutoff (inches) | 0.2
LCD Totalizer Multiplier
Go to: Setup/Totalizers (2.16) |  | 10
Totalizer Decimal (0,1,2,3) | 0
Totalizer Multiplier | 10 | If flow units GPM, multiply displayed gallons
by this number
To assign relay for remote totalizer,
Go to: Setup/Other Control/External Totalizer |  | If rate is MGD, .001 makes totalizer read
Enable | No | Gallons X 1000, (or million gallons X.001)
Multiplier | 100
Relay duration (sec) | 0.2
Assigned Relay (1,2,3)
Cycle Power for all parameters to take effect
Remote Totalizer Multiplier
100
Enter the factor by which the volume is
divided, prior to remote totalizer actuation
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421

## Rectangular weir tables (hidden)
Size | Multiplier | Head (ft) |  |  | Sensor type
1' Rectangular weir with end contractions | 1 | 0.039166 | 1 | 1 | STH | 100
1.5' Rectangular weir with end contractions | 1.5 | 0.025044 | 1 |  | XPS-10 | 102 | 4
2' Rectangular weir with end contractions | 2 | 0.019904 | 1 |  | XPS-15 | 104
2.5' Rectangular weir with end contractions | 2.5 | 0.013269 | 1 |  | XRS-5 | 112
3' Rectangular weir with end contractions | 3 | 0.0088757 | 1 |  | Milliamp input | 250
4' Rectangular weir with end contractions | 4 | 0.0072775 | 1
5' Rectangular weir with end contractions | 5 | 0.0058841 | 1 |  | Flow units | Multiplier | Time
6' Rectangular weir with end contractions | 6 | 0.0050862 | 1 |  | GPM | 1 | 2 | Maximum flow rate (GPM) | 2
8' Rectangular weir with end contractions | 8 | 0.004544 | 1 |  | MGD | 0.00144 | 4 | Maximum flow rate (MGD)
10' Rectangular weir with end contractions | 10 | 0.0041431 | 1
Totalizer scale factor
[5 empty] | 0.001 | -3 | 6
Graph of Y= 2.152*((Width)-0.2*(x))*X^1.5 |  |  |  |  | 0.01 | -2
X | Y |  |  |  | 0.1 | -1
1.0002 | 0.72203 |  |  |  | 1 | 0
[5 empty] | 10 | 1
[5 empty] | 100 | 2
[5 empty] | 1000 | 3
[5 empty] | 10000 | 4
[5 empty] | 100000 | 5
[5 empty] | 1000000 | 6
[5 empty] | 10000000 | 7
Totalizer Decimal Position
[5 empty] | No decimal | 0 | 3
[5 empty] | 1 digit after decimal | 1
[5 empty] | 2 digits after decimal | 2
[5 empty] | 3 digits after decimal | 3
Primary Display
[5 empty] | Flow rate | 1 | 1
[5 empty] | Flow total | 2
[5 empty] | Toggle flow/total, 3 seconds each | 3
Remote totalizer
[5 empty] | Remote totalizer on | 40 | 1
[5 empty] | Remote totalizer off | 0
Remote totalizer scale factor
[5 empty] | 0.001 | -3 | 3
[5 empty] | 0.01 | -2
[5 empty] | 0.1 | -1
[5 empty] | 1 | 0
[5 empty] | 10 | 1
[5 empty] | 100 | 2
[5 empty] | 1000 | 3
[5 empty] | 10000 | 4
[5 empty] | 100000 | 5
[5 empty] | 1000000 | 6
[5 empty] | 10000000 | 7
Rate of Change
[5 empty] | Rate of Change: Slow | 1 | 1 | Rate of Change: Slow
[5 empty] | Rate of Change: Medium | 2 |  | Rate of Change: Medium
[5 empty] | Rate of Change: Fast | 3 |  | Rate of Change: Fast

## LUT400 Rect weir end contract.
LUT400 Open Channel Flowmeter
Sensor Type
XPS-15
Date:
Location:
Primary Device 3' Rectangular weir with end contractions
Max flow (MGD): | 2
Rate of Change
Go to: Wizards/Quick Start/QS Flow |  | Rate of Change: Medium
Description | Value
Transducer | XPS-10
Temperature Source | Transducer
Primary Measuring Device | Rect weir contracted | Primary Device
Method of Flow Calculation | Ratiometric | 3' Rectangular weir with end contractions
Units | Inches
High Calibration Point | -5.6128
Low Calibration Point (sensor face to zero elevation)
Response Rate (slow, med, fast) | Medium
Crest width | 36
Maximum Head | 5.6128
Zero Head offset | 0
Flow Rate Units | MGD (MMGAL/D)
Maximum flow at 20 mA (MGD) | 2
Flow Rate Decimal (0,1,2,3) | 0
Low Flow Cutoff (inches) | 0.2
LCD Totalizer Multiplier
Go to: Setup/Totalizers (2.16) |  | 0.001
Totalizer Decimal (0,1,2,3) | 0
Totalizer Multiplier | 0.001 | If flow units GPM, multiply displayed gallons
by this number
To assign relay for remote totalizer,
Go to: Setup/Other Control/External Totalizer |  | If rate is MGD, .001 makes totalizer read
Enable | No | Gallons X 1000, (or million gallons X.001)
Multiplier | 10
Relay duration (sec) | 0.2
Assigned Relay (1,2,3)
Cycle Power for all parameters to take effect
Remote Totalizer Multiplier
10
Enter the factor by which the volume is
divided, prior to remote totalizer actuation
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421
Click to calculate max head

## LUT400 Rect weir tables (hidden)
Size | Multiplier | Head (ft) |  |  | Sensor type
1' Rectangular weir with end contractions | 1 | 0 | 1 | 5 | STH | 100
1.5' Rectangular weir with end contractions | 1.5 | 0 | 1 |  | XPS-10 | 102 | 2
2' Rectangular weir with end contractions | 2 | 0 | 1 |  | XPS-15 | 104
2.5' Rectangular weir with end contractions | 2.5 | 0 | 1 |  | XRS-5 | 112
3' Rectangular weir with end contractions | 3 | 0 | 1 |  | Milliamp input | 250
4' Rectangular weir with end contractions | 4 | 0 | 1
5' Rectangular weir with end contractions | 5 | 0 | 1 |  | Flow units | Multiplier | Time
6' Rectangular weir with end contractions | 6 | 0 | 1 |  | GPM | 1 | 2 | Maximum flow rate (GPM) | 1
8' Rectangular weir with end contractions | 8 | 0 | 1 |  | MGD | 0.00144 | 4 | Maximum flow rate (MGD)
10' Rectangular weir with end contractions | 10 | 0 | 1
Totalizer scale factor
[5 empty] | 0.001 | -3 | 1
Graph of Y= 2.152*((Width)-0.2*(x))*X^1.5 |  |  |  |  | 0.01 | -2
X | Y |  |  |  | 0.1 | -1
0.46773 | 0.00080459 |  |  |  | 1 | 0
[5 empty] | 10 | 1
[5 empty] | 100 | 2
Primary Device: 1' Rectangular weir with end contractions |  |  |  |  | 1000 | 3
Primary Device: 1.5' Rectangular weir with end contractions |  |  |  |  | 10000 | 4
Primary Device: 2' Rectangular weir with end contractions |  |  |  |  | 100000 | 5
Primary Device 2.5' Rectangular weir with end contractions |  |  |  |  | 1000000 | 6
Primary Device 3' Rectangular weir with end contractions |  |  |  |  | 10000000 | 7
Primary Device 4' Rectangular weir with end contractions
Primary Device 5' Rectangular weir with end contractions |  |  |  |  | Totalizer Decimal Position
Primary Device 6' Rectangular weir with end contractions |  |  |  |  | No decimal | 0 | 3
Primary Device: 8' Rectangular weir with end contractions |  |  |  |  | 1 digit after decimal | 1
Primary Device: 10' Rectangular weir with end contractions |  |  |  |  | 2 digits after decimal | 2
[5 empty] | 3 digits after decimal | 3
Primary Display
[5 empty] | Flow rate | 1 | 3
[5 empty] | Flow total | 2
[5 empty] | Toggle flow/total, 3 seconds each | 3
Remote totalizer
[5 empty] | Remote totalizer on | 40 | 1
[5 empty] | Remote totalizer off | 0
Remote totalizer scale factor
[5 empty] | 0.001 | -3 | 5
[5 empty] | 0.01 | -2
[5 empty] | 0.1 | -1
[5 empty] | 1 | 0
[5 empty] | 10 | 1
[5 empty] | 100 | 2
[5 empty] | 1000 | 3
[5 empty] | 10000 | 4
[5 empty] | 100000 | 5
[5 empty] | 1000000 | 6
[5 empty] | 10000000 | 7
Rate of Change
[5 empty] | Rate of Change: Slow | 1 | 2 | Rate of Change: Slow
[5 empty] | Rate of Change: Medium | 2 |  | Rate of Change: Medium
[5 empty] | Rate of Change: Fast | 3 |  | Rate of Change: Fast

## LT500 OCF files (hidden)
Size | Exponent | Head (in) | [8 empty] | Sensor type
1" Parshall flume | 1.55 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head | 4 | LR110 | LR110 | 3
2" Parshall flume | 1.55 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | LR120 | LR120
3" Parshall Flume | 1.55 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | LU240 | LU240
6" Parshall Flume | 1.58 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Generic 4-20 mA | 4-20 mA
9" Parshall flume | 1.53 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
12" Parshall flume | 1.522 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
18" Parshall flume | 1.538 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Flow units | Multiplier | Time
24" Parshall flume | 1.55 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | GPM | 1 | 2 | Maximum flow rate (GPM) | 2
30" Parshall flume | 1.559 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Mgal/d | 0.00144 | 4 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  |  |  | 0.00144
4' Parshall flume | 1.578 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
5' Parshall flume | 1.587 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Totalizer units
6' Parshall flume | 1.595 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Gallon | 264.2 | 3
8' Parshall flume | 1.607 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Kgal | 0.2642
10' Parshall flume | 1.6 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Mgal | 0.0002642
12' Parshall flume | 1.6 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Cubic Feet | 35.134
4" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
6" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
8" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
10" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Totalizer Decimal Position
12" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | None | None | 3
15" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 1 digit after decimal | 1
18" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 2 digits after decimal | 2
21" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 3 digits after decimal | 3
24" Leopold Lagco flume | 1.547 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 4 digits after decimal | 4
Trapezoidal, small 60 deg | 2.58 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 5 digits after decimal | 5
Trapezoidal, large 60 deg V | 2.58 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | 6 digits after decimal | 6
Trapezoidal, XL 60 degree V | 2.6363 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
Trapezoidal, 3' 60 degree V | 2.5804 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
V-notch, 22.5 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
V-notch, 30 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
V-notch, 45 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Primary Display
V-notch, 60 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Flow rate | 1 | 2
V-notch, 90 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Flow total | 2
V-notch, 120 degree | 2.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head |  | Toggle flow/total, 3 seconds each | 3
1' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
1.5' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
2' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
2.5' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
3' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
4' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
5' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
6' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
8' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
10' Rectangular weir without end contractions | 1.5 | 0 | 0 | Exponential | Ratiometric | Method of Flow Calculation | Flow Exponent | Upper calibration point | Maximum head
1' Rectangular Weir, Contracted | 12" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
1 1/2' Rectangular Weir, Contracted | 18" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
2' Rectangular Weir, Contracted | 24" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
2.5' Rectangular Weir, Contracted | 30" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
3' Rectangular Weir, Contracted | 36" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
4' Rectangular Weir, Contracted | 48" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
5' Rectangular Weir, Contracted | 60" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
6' Rectangular Weir, Contracted | 72" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
8' Rectangular Weir, Contracted | 96" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
10' Rectangular Weir, Contracted | 120" |  |  | Rectangular weir contracted | Absolute | Method of Flow Calculation | Crest width | Upper calibration point (see tables) | Maximum head (see tables)
4" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 4 | Maximum Flume Width |  | Upper calibration point | Maximum head
6" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 6 | Maximum Flume Width |  | Upper calibration point | Maximum head
8" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 8 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | Rate of Change |  |  |  | View 2 (enabled/disabled
10" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 10 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | Slow | Slow | 3 |  | Enabled | 1
12" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 12 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | Medium | Medium |  |  | Disabled
15" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 15 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | Fast | Fast
18" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 18 | Maximum Flume Width |  | Upper calibration point | Maximum head
21" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 21 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | View 1 |  |  |  | View 2
24" Palmer Bowlus Flume |  | 0 | 0 | Palmer Bowlus | 27 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | Type |  |  |  | Type
27" Palmer Bowlus Flume |  | -0.468 | 0.468 | Palmer Bowlus | 27 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 1 value |  | 6 |  | 1 value | 5
30" Palmer Bowlus Flume |  | -0.75 | 0.75 | Palmer Bowlus | 30 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 1 Value and horizontal bargraph |  |  |  | 1 Value and horizontal bargraph
36" Palmer Bowlus Flume |  | -0.96 | 0.96 | Palmer Bowlus | 36 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 1 value, icons and vertical bargraph |  |  |  | 1 value, icons and vertical bargraph
42" Palmer Bowlus Flume |  | -0.936 | 0.936 | Palmer Bowlus | 42 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 1 value and trend chart |  |  |  | 1 value and trend chart
48" Palmer Bowlus Flume |  | -1.02 | 1.02 | Palmer Bowlus | 48 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 3 values and horizontal bargraph |  |  |  | 3 values and horizontal bargraph
60" Palmer Bowlus Flume |  | -1.512 | 1.512 | Palmer Bowlus | 60 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 3 values, icons and vertical bargraph |  |  |  | 3 values, icons and vertical bargraph
72" Palmer Bowlus Flume |  | -1.86 | 1.86 | Palmer Bowlus | 72 | Maximum Flume Width |  | Upper calibration point | Maximum head |  | 6 values |  |  |  | 6 values
0.5' H Flume |  |  |  | H-Flume | 6" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables)
0.75' H Flume |  |  |  | H-Flume | 9" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables)
1.0' H Flume |  |  |  | H-Flume | 12" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | 1st value |  |  |  | 1st value
1.5' H Flume |  |  |  | H-Flume | 18" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | Head (point 1) |  | 2 |  | Head (point 1) | 6
2.0' H Flume |  |  |  | H-Flume | 24" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | Volume flow (point 1) |  |  |  | Volume flow (point 1)
2.5' H Flume |  |  |  | H-Flume | 30" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | Volume flow% (point 1) |  |  |  | Volume flow% (point 1)
3.0' H Flume |  |  |  | H-Flume | 36" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | Sensor temperature |  |  |  | Sensor temperature
4.5' H Flume |  |  |  | H-Flume | 54" | Flume Height |  | Upper calibration point (see tables) | Maximum head (see tables) |  | Totalizer 1 |  |  |  | Totalizer 1
[11 empty] | Totalizer 2 |  |  |  | Totalizer 2
[11 empty] | Totalizer 3 |  |  |  | Totalizer 3
[11 empty] | Totalizer 4 |  |  |  | Totalizer 4
[11 empty] | 2nd value |  |  |  | 2nd value
[11 empty] | None |  |  |  | None
[11 empty] | Head (point 1) |  | 2 |  | Head (point 1) | 7
[11 empty] | Volume flow (point 1) |  |  |  | Volume flow (point 1)
[11 empty] | Volume flow% (point 1) |  |  |  | Volume flow% (point 1)
[11 empty] | Sensor temperature |  |  |  | Sensor temperature
[11 empty] | Totalizer 1 |  |  |  | Totalizer 1
[11 empty] | Totalizer 2 |  |  |  | Totalizer 2
[11 empty] | Totalizer 3 |  |  |  | Totalizer 3
[11 empty] | Totalizer 4 |  |  |  | Totalizer 4
[11 empty] | 3rd value |  |  |  | 3rd value
[11 empty] | None |  |  |  | None
[11 empty] | Head (point 1) |  | 6 |  | Head (point 1) | 8
[11 empty] | Volume flow (point 1) |  |  |  | Volume flow (point 1)
[11 empty] | Volume flow% (point 1) |  |  |  | Volume flow% (point 1)
[11 empty] | Sensor temperature |  |  |  | Sensor temperature
[11 empty] | Totalizer 1 |  |  |  | Totalizer 1
[11 empty] | Totalizer 2 |  |  |  | Totalizer 2
[11 empty] | Totalizer 3 |  |  |  | Totalizer 3
[11 empty] | Totalizer 4 |  |  |  | Totalizer 4

## LR100 OCF
LR100 Open Channel Flowmeter
Date:
Location:
Bluetooth PIN
Primary Device:
Max flow (GPM):
Temperature Units | F
Level Units | in
Operation | Volume Flow
Material Type | Liquid
Application | Flume/weir
Lower Calibration Point
Upper Calibration Point
Volume Flow Units
Primary Measuring Device
Range (upper range value)

## Probe-LU-LR OCF
Siemens Probe-LU Open Channel Flowmeter |  |  | Rate of Change
Date:
Location: |  |  | Primary Device
6" Parshall Flume
Primary Device: |  | 6" Parshall Flume
Max. flow (GPM):
Parameter | Value | Description
P001 | 1 | Operation: Level
P003 | 2 | Rate of change: Medium | Flow Units
P005 | 5 | Units of Measure (Inches) | GPM
P006 |  | *Distance from sensor face zero flow level
P007 | 0 | Head at maximum flow (inches)
P050 | 9 | Vessel Shape: open channel flow breakpoints
P051 | 0 | Maximum flow rate (GPM)
P054, 01 | 0 | Level #1 | P055, 01 | 0 | Flow #1
P054, 02 | 0.19 | Level #2 | P055, 02 | 0.001 | Flow #2
P054, 03 | 0.2 | Level #3 (low flow cutoff= 0.2") | P055, 03 | 1.4336 | Flow #3
P054, 04 | 0 | Level #4 | P055, 04 | 0 | Flow #4
P054, 05 | 0 | Level #5 | P055, 05 | 0 | Flow #5
P054, 06 | 0 | Level #6 | P055, 06 | 0 | Flow #6
P054, 07 | 0 | Level #7 | P055, 07 | 0 | Flow #7
P054, 08 | 0 | Level #8 | P055, 08 | 0 | Flow #8
P054, 09 | 0 | Level #9 | P055, 09 | 0 | Flow #9
P054, 10 | 0 | Level #10 | P055, 10 | 0 | Flow #10
P054, 11 | 0 | Level #11 | P055, 11 | 0 | Flow #11
P054, 12 | 0 | Level #12 | P055, 12 | 0 | Flow #12
P054, 13 | 0 | Level #13 | P055, 13 | 0 | Flow #13
P054, 14 | 0 | Level #14 | P055, 14 | 0 | Flow #14
P054, 15 | 0 | Level #15 | P055, 15 | 0 | Flow #15
P054, 16 | 0 | Level #16 | P055, 16 | 0 | Flow #16
P054, 17 | 0 | Level #17 | P055, 17 | 0 | Flow #17
P054, 18 | 0 | Level #18 | P055, 18 | 0 | Flow #18
P054, 19 | 0 | Level #19 | P055, 19 | 0 | Flow #19
P054, 20 | 0 | Level #20 | P055, 20 | 0 | Flow #20
P054, 21 | 0 | Level #21 | P055, 21 | 0 | Flow #21
P054, 22 | 0 | Level #22 | P055, 22 | 0 | Flow #22
P054, 23 | 0 | Level #23 | P055, 23 | 0 | Flow #23
P054, 24 | 0 | Level #24 | P055, 24 | 0 | Flow #24
P054, 25 | 0 | Level #25 | P055, 25 | 0 | Flow #25
P054, 26 | 0 | Level #26 | P055, 26 | 0 | Flow #26
P054, 27 | 0 | Level #27 | P055, 27 | 0 | Flow #27
P054, 28 | 0 | Level #28 | P055, 28 | 0 | Flow #28
P054, 29 | 0 | Level #29 | P055, 29 | 0 | Flow #29
P054, 30 | 0 | Level #30 | P055, 30 | 0 | Flow #30
P054, 31 | 0 | Level #31 | P055, 31 | 0 | Flow #31
P054, 32 | 0 | Level #32 | P055, 32 | 0 | Flow #32
P201 | 4 | Analog output function (flow rate)
P801 |  | * Range extension (inches)
* Measured variables
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Level | Rate
inches | GPM
0 | 0
0.19 | 0.001
0.2 | 1.4336
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
0 | 0
Note on D50: The distance below zero flow condition (PO06) that the level can fall to. This allows water level to fall below 'zero flow' level, without reporting Loss of Echo (LOE).

## Probe LU tables (hidden)
Size | Exponent | K-factor | Head (ft)
1" Parshall flume | 1.55 | 151.7 | 0 | 0 | 4
2" Parshall flume | 1.55 | 303.4 | 0 | 0
3" Parshall Flume | 1.557 | 445.2 | 0 | 0
6" Parshall Flume | 1.58 | 924.5 | 0 | 0
9" Parshall flume | 1.53 | 1378 | 0 | 0
12" Parshall flume | 1.522 | 1795 | 0 | 0
18" Parshall flume | 1.538 | 2693 | 0 | 0 |  | Flow units | Multiplier
24" Parshall flume | 1.55 | 3590 | 0 | 0 |  | GPM | 1 | Maximum flow rate (GPM) | 1
30" Parshall flume | 1.559 | 4488 | 0 | 0 |  | MGD | 0.00144 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 5386 | 0 | 0
4'Parshall flume | 1.578 | 7181 | 0 | 0
5' Parshall flume | 1.587 | 8976 | 0 | 0 |  | Rate of Change
6' Parshall flume | 1.595 | 10770 | 0 | 0 |  | Rate of change: Slow | 1 | 2 | Rate of change: Slow
8' Parshall flume | 1.607 | 14360 | 0 | 0 |  | Rate of change: Medium | 2 |  | Rate of change: Medium
10' Parshall flume | 1.6 | 17670 | 0 | 0 |  | Rate of change: Fast | 3 |  | Rate of change: Fast
12' Parshall flume | 1.55 | 20980 | 0 | 0
4" Leopold Lagco flume | 1.547 | 379.1 | 0 | 0
6" Leopold Lagco flume | 1.547 | 557.9 | 0 | 0
8" Leopold Lagco flume | 1.547 | 733.9 | 0 | 0
10" Leopold Lagco flume | 1.547 | 907.7 | 0 | 0
12" Leopold Lagco flume | 1.547 | 1080 | 0 | 0
15" Leopold Lagco flume | 1.547 | 1336 | 0 | 0
18" Leopold Lagco flume | 1.547 | 1589 | 0 | 0
21" Leopold Lagco flume | 1.547 | 1841 | 0 | 0
24" Leopold Lagco flume | 1.547 | 2091 | 0 | 0
Trapezoidal, small 60 deg | 2.58 | 695.6 | 0 | 0
Trapezoidal, large 60 deg V | 2.58 | 695.6 | 0 | 0
Trapezoidal, XL 60 degree V | 2.6363 | 694.7 | 0 | 0
Trapezoidal, 3' 60 degree V | 2.5804 | 594.6 | 0 | 0
V-notch, 22.5 degree | 2.5 | 223.1 | 0 | 0
V-notch, 30 degree | 2.5 | 303.4 | 0 | 0
V-notch, 45 degree | 2.5 | 464.5 | 0 | 0
V-notch, 60 degree | 2.5 | 647.6 | 0 | 0
V-notch, 90 degree | 2.5 | 1122 | 0 | 0
V-notch, 120 degree | 2.5 | 1943 | 0 | 0
1' Rectangular weir without end contractions | 1.5 | 1495 | 0 | 0
1.5' Rectangular weir without end contractions | 1.5 | 2242 | 0 | 0
2' Rectangular weir without end contractions | 1.5 | 2989 | 0 | 0
2.5' Rectangular weir without end contractions | 1.5 | 3736 | 0 | 0
3' Rectangular weir without end contractions | 1.5 | 4484 | 0 | 0
4' Rectangular weir without end contractions | 1.5 | 5978 | 0 | 0
5' Rectangular weir without end contractions | 1.5 | 7473 | 0 | 0
6' Rectangular weir without end contractions | 1.5 | 8967 | 0 | 0
8' Rectangular weir without end contractions | 1.5 | 11960 | 0 | 0
10' Rectangular weir without end contractions | 1.5 | 14950 | 0 | 0
4" Palmer Bowlus Flume | 1.9573 | 776.4 | -0.00588 | 0.00588
6" Palmer Bowlus Flume | 1.9025 | 929.4 | -0.005421 | 0.005421
8" Palmer Bowlus Flume | 1.9724 | 1138.6 | -0.01456 | 0.01456
10" Palmer Bowlus Flume | 1.953 | 1275.9 | -0.0161 | 0.0161
12" Palmer Bowlus Flume | 1.9362 | 1410.9 | -0.017 | 0.017
15" Palmer Bowlus Flume | 1.9062 | 1604 | -0.01682 | 0.01682
18" Palmer Bowlus Flume | 1.8977 | 1789.8 | -0.01875 | 0.01875
21" Palmer Bowlus Flume | 1.9619 | 1895 | -0.039 | 0.039
27" Palmer Bowlus Flume | 1.9269 | 2230 | -0.038 | 0.038
30" Palmer Bowlus Flume | 1.9663 | 2254 | -0.0625 | 0.0625
36" Palmer Bowlus Flume | 1.991 | 2451 | -0.08 | 0.08
42" Palmer Bowlus Flume | 1.9268 | 2747 | -0.078 | 0.078
48" Palmer Bowlus Flume | 1.9586 | 2974 | -0.085 | 0.085
60" Palmer Bowlus Flume | 1.9833 | 3224 | -0.126 | 0.126
72" Palmer Bowlus Flume | 1.9871 | 3518 | -0.155 | 0.155

## LR200_LR250 OCF
Siemens LR200/LR250 Open Channel Flowmeter
Primary Device
Date: |  |  | 12' Parshall flume
Location:
Primary Device: |  | 12' Parshall flume
Max. flow (GPM): |  | 138888
Parameter | Value | Description
Quick Start
Material | Liquid | Operation: Level
Response | Medium | Rate of change: Medium | Flow Units
Units | Inches | Units of Measure (Inches) | GPM
Operation | Level
Low cal point | 91.32 | *Distance from reference point at zero flow
Hi cal point | 52.2163 | Distance from reference point at max. flow
2.6.1 | Volume | Current output = flow rate | Rate of Change
2.7.1.1 | Linear Table | Level converted to flow rate via table | Rate of change: Medium
2.7.1.2 | 138888 | Maximum flow rate (GPM)
2.8.2 | 112 | * Far range
* = Measured variable
Level to Flow Data |  | (go to: Setup/Linearization/Table)
2.7.2.1 | 0 | Level #1 (all levels in inches) | 2.7.2.2 | 0 | Flow #1
2.7.2.3 | 0.19 | Level #2 | 2.7.2.4 | 0.001 | Flow #2
2.7.2.5 | 0.2 | Level #3 (low flow cutoff= 0.2") | 2.7.2.6 | 29.9753 | Flow #3
2.7.2.7 | 1.9552 | Level #4 | 2.7.2.8 | 1150.8447 | Flow #4
2.7.2.9 | 3.9104 | Level #5 | 2.7.2.10 | 3488.7088 | Flow #5
2.7.2.11 | 5.8656 | Level #6 | 2.7.2.12 | 6674.3771 | Flow #6
2.7.2.13 | 7.8207 | Level #7 | 2.7.2.14 | 10575.7875 | Flow #7
2.7.2.15 | 9.7759 | Level #8 | 2.7.2.16 | 15113.6283 | Flow #8
2.7.3.1 | 11.7311 | Level #9 | 2.7.3.2 | 20232.9278 | Flow #9
2.7.3.3 | 13.2953 | Level #10 | 2.7.3.4 | 24719.0053 | Flow #10
2.7.3.5 | 14.4684 | Level #11 | 2.7.3.6 | 28300.0682 | Flow #11
2.7.3.7 | 15.6415 | Level #12 | 2.7.3.8 | 32059.7927 | Flow #12
2.7.3.9 | 16.8146 | Level #13 | 2.7.3.10 | 35992.6856 | Flow #13
2.7.3.11 | 17.9877 | Level #14 | 2.7.3.12 | 40093.8032 | Flow #14
2.7.3.13 | 19.1608 | Level #15 | 2.7.3.14 | 44358.6634 | Flow #15
2.7.3.15 | 20.3339 | Level #16 | 2.7.3.16 | 48783.1763 | Flow #16
2.7.4.1 | 21.5071 | Level #17 | 2.7.4.2 | 53363.5886 | Flow #17
2.7.4.3 | 22.6802 | Level #18 | 2.7.4.4 | 58096.4391 | Flow #18
2.7.4.5 | 23.8533 | Level #19 | 2.7.4.6 | 62978.5214 | Flow #19
2.7.4.7 | 25.0264 | Level #20 | 2.7.4.8 | 68006.8536 | Flow #20
2.7.4.9 | 26.1995 | Level #21 | 2.7.4.10 | 73178.6525 | Flow #21
2.7.4.11 | 27.3726 | Level #22 | 2.7.4.12 | 78491.3119 | Flow #22
2.7.4.13 | 28.5457 | Level #23 | 2.7.4.14 | 83942.3842 | Flow #23
2.7.4.15 | 29.7188 | Level #24 | 2.7.4.16 | 89529.5645 | Flow #24
2.7.5.1 | 30.892 | Level #25 | 2.7.5.2 | 95250.6768 | Flow #25
2.7.5.3 | 32.0651 | Level #26 | 2.7.5.4 | 101103.6624 | Flow #26
2.7.5.5 | 33.2382 | Level #27 | 2.7.5.6 | 107086.5692 | Flow #27
2.7.5.7 | 34.4113 | Level #28 | 2.7.5.8 | 113197.5428 | Flow #28
2.7.5.9 | 35.5844 | Level #29 | 2.7.5.10 | 119434.8179 | Flow #29
2.7.5.11 | 36.7575 | Level #30 | 2.7.5.12 | 125796.712 | Flow #30
2.7.5.13 | 37.9306 | Level #31 | 2.7.5.14 | 132281.6182 | Flow #31
2.7.5.15 | 39.1037 | Level #32 | 2.7.5.16 | 138888 | Flow #32
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Level | Rate
inches | GPM
0 | 0
0.19 | 0.001
0.2 | 29.9753
1.9552 | 1150.8447
3.9104 | 3488.7088
5.8656 | 6674.3771
7.8207 | 10575.7875
9.7759 | 15113.6283
11.7311 | 20232.9278
13.2953 | 24719.0053
14.4684 | 28300.0682
15.6415 | 32059.7927
16.8146 | 35992.6856
17.9877 | 40093.8032
19.1608 | 44358.6634
20.3339 | 48783.1763
21.5071 | 53363.5886
22.6802 | 58096.4391
23.8533 | 62978.5214
25.0264 | 68006.8536
26.1995 | 73178.6525
27.3726 | 78491.3119
28.5457 | 83942.3842
29.7188 | 89529.5645
30.892 | 95250.6768
32.0651 | 101103.6624
33.2382 | 107086.5692
34.4113 | 113197.5428
35.5844 | 119434.8179
36.7575 | 125796.712
37.9306 | 132281.6182
39.1037 | 138888
Note on D20: Distance below reference point that level can fall to and still not register loss of echo. Useful if level can drop below low calibration point.

## LR200-250 tables (hidden)
Size | Exponent | K-factor | Head (ft)
1" Parshall flume | 1.55 | 151.7 | 81.4256 | 0 | 16 | Rate of Change
2" Parshall flume | 1.55 | 303.4 | 52.0652 | 0 |  | Slow | 1 | 2 | Rate of change: Slow
3" Parshall Flume | 1.557 | 445.2 | 39.9824 | 0 |  | Medium | 2 |  | Rate of change: Medium
6" Parshall Flume | 1.58 | 924.5 | 23.8613 | 0 |  | Fast | 3 |  | rate of change: Fast
9" Parshall flume | 1.53 | 1378 | 20.3902 | 0
12" Parshall flume | 1.522 | 1795 | 17.4128 | 0
18" Parshall flume | 1.538 | 2693 | 12.9841 | 0 |  | Flow units | Multiplier
24" Parshall flume | 1.55 | 3590 | 10.574 | 0 |  | GPM | 1 | Maximum flow rate (GPM) | 1
30" Parshall flume | 1.559 | 4488 | 9.0393 | 0 |  | MGD | 0.00144 | Maximum flow rate (MGD)
3' Parshall flume | 1.566 | 5386 | 7.9667 | 0
4'Parshall flume | 1.578 | 7181 | 6.5352 | 0
5' Parshall flume | 1.587 | 8976 | 5.618 | 0
6' Parshall flume | 1.595 | 10770 | 4.9683 | 0 |  |  | LR200_LR250 OCF meter
8' Parshall flume | 1.607 | 14360 | 4.1045 | 0
10' Parshall flume | 1.6 | 17670 | 3.6278 | 0
12' Parshall flume | 1.6 | 20980 | 3.2586 | 0
4" Leopold Lagco flume | 1.547 | 379.1 | 45.4301 | 0
6" Leopold Lagco flume | 1.547 | 557.9 | 35.3895 | 0
8" Leopold Lagco flume | 1.547 | 733.9 | 29.6414 | 0
10" Leopold Lagco flume | 1.547 | 907.7 | 25.8364 | 0
12" Leopold Lagco flume | 1.547 | 1080 | 23.0908 | 0
15" Leopold Lagco flume | 1.547 | 1336 | 20.0078 | 0
18" Leopold Lagco flume | 1.547 | 1589 | 17.9902 | 0
21" Leopold Lagco flume | 1.547 | 1841 | 16.3573 | 0
24" Leopold Lagco flume | 1.547 | 2091 | 15.0648 | 0
Trapezoidal, small 60 deg | 2.58 | 695.6 | 7.791 | 0
Trapezoidal, large 60 deg V | 2.58 | 695.6 | 7.791 | 0
Trapezoidal, XL 60 degree V | 2.6363 | 694.7 | 7.7949 | 0
Trapezoidal, 3' 60 degree V | 2.5804 | 594.6 | 8.2767 | 0
V-notch, 22.5 degree | 2.5 | 223.1 | 13.1119 | 0
V-notch, 30 degree | 2.5 | 303.4 | 11.5947 | 0
V-notch, 45 degree | 2.5 | 464.5 | 9.7785 | 0
V-notch, 60 degree | 2.5 | 647.6 | 8.5614 | 0
V-notch, 90 degree | 2.5 | 1122 | 6.8718 | 0
V-notch, 120 degree | 2.5 | 1943 | 5.5167 | 0
1' Rectangular weir without end contractions | 1.5 | 1495 | 20.5124 | 0
1.5' Rectangular weir without end contractions | 1.5 | 2242 | 15.6562 | 0
2' Rectangular weir without end contractions | 1.5 | 2989 | 12.9249 | 0
2.5' Rectangular weir without end contractions | 1.5 | 3736 | 11.1388 | 0
3' Rectangular weir without end contractions | 1.5 | 4484 | 9.8628 | 0
4' Rectangular weir without end contractions | 1.5 | 5978 | 8.1422 | 0
5' Rectangular weir without end contractions | 1.5 | 7473 | 7.0164 | 0
6' Rectangular weir without end contractions | 1.5 | 8967 | 6.2136 | 0
8' Rectangular weir without end contractions | 1.5 | 11960 | 5.1281 | 0
10' Rectangular weir without end contractions | 1.5 | 14950 | 4.4193 | 0
4" Palmer Bowlus Flume | 1.9573 | 776.4 | 14.1475 | 0.00588
6" Palmer Bowlus Flume | 1.9025 | 929.4 | 13.8925 | 0.005421
8" Palmer Bowlus Flume | 1.9724 | 1138.6 | 11.4075 | 0.01456
10" Palmer Bowlus Flume | 1.953 | 1275.9 | 11.023 | 0.0161
12" Palmer Bowlus Flume | 1.9362 | 1410.9 | 10.684 | 0.017
15" Palmer Bowlus Flume | 1.9062 | 1604 | 10.368 | 0.01682
18" Palmer Bowlus Flume | 1.8977 | 1789.8 | 9.8866 | 0.01875
21" Palmer Bowlus Flume | 1.9619 | 1895 | 8.8866 | 0.039
27" Palmer Bowlus Flume | 1.9269 | 2230 | 8.4972 | 0.038
30" Palmer Bowlus Flume | 1.9663 | 2254 | 8.0694 | 0.0625
36" Palmer Bowlus Flume | 1.991 | 2451 | 7.5167 | 0.08
42" Palmer Bowlus Flume | 1.9268 | 2747 | 7.3019 | 0.078
48" Palmer Bowlus Flume | 1.9586 | 2974 | 7.0321 | 0.085
60" Palmer Bowlus Flume | 1.9833 | 3224 | 6.5423 | 0.126
72" Palmer Bowlus Flume | 1.9871 | 3518 | 6.2037 | 0.155

## OCM-III OCF
Siemens OCM-III Open Channel Flowmeter
Primary Device
Date: |  |  | 18" Parshall flume
Location:
Primary Device: |  | 18" Parshall flume
Max. flow (GPM): |  | 694
Parameter | Value | Description
P0 | 0 | Language: English | Flow Units
P1 | 1 | Units of Measure (Inches) | GPM
P2 | 1 | Temperature units
P3 | 0 | Type of primary device
P4 | 1 | Method of calculation
P5 | 3 | Flow rate units | Totalizer Multiplier
P6 | 694 | Maximum flow rate (GPM) | 100
P7 | 4.9693 | Head at max flow (inches) | If flow units GPM, multiply gallons by this number
P24 | 0 | Ma assignment: Flow rate | If flow units MGD, multiply million gal. by this number
P26 | 0 | Output range (4-20 mA)
P32 | 5 | LCD Totalizer multiplier (Reading x multiplier)
P33 | 2 | Flow rate display decimal position
P42 | 0 | Head determination: by OCM-III
P45 | 0.2 | Low flow cutoff head
P46 |  | *Range at Zero head
U0 | 1.538 | Exponent value,
|  | Flume width if Palmer Bowlus Flume, | Flow Rate Decimal Position
|  | or crest width if contracted rectangular weir | 2 digits after decimal
P15 | 0 | Remote totalizer from relay 1
P16 | 1 | Remote totalizer acts with LCD totalizer
* Measured variable
Notes: |  |  | Temperature Units
1: Enter 'F2' to save program, and return to RUN mode |  |  | Deg F
Password: 2.71828
Remote totalizer from relay 1
no
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Note on D21: Example: Multiplier = 100 Totalizer display = 35 Actual total = 3500
Note on D24: Level (in inches) below which totalizer activity is stopped
Note on D25: Distance from sensor face to water level at zero flow.

## Well pump control
[5 empty] | Rate of change | Relay 1 control function
Siemens HydroRanger 200, Multiranger 200, and Enviroranger Pump Control
Location: |  |  |  |  | Sensor type
Date: | [5 empty] | Relay 2 control function
Parameter | Value | Description |  |  | Units
P001 | 1 | Operation (1= level) |  |  |  | Relay 3 control function
P002 | 1 | Material (1= liquid)
P003 | 2 | Maximum process speed
P004 | 104 | Sensor type |  |  | Failsafe material level
P005 | 5 | Units |  |  |  | Relay 4 control function
P006 |  | Distance from sensor face to min level
P007 |  | Maximum level above zero level
P064 | 0 | Reading override enable |  |  | Output failsafe
P065 |  | Reading override value |  |  |  | Relay 5 control function
P071 | HOLD | Failsafe material level
P110 | 1 | Level source (1= transducer 1)
P111(1) | 0 | Relay #1 control function |  |  | Auxiliary high or low level switch
P111(2) | 0 | Relay #2 control function |  |  |  | Relay 6 control function
P111(3) | 0 | Relay #3 control function
P111(4) | 0 | Relay #4 control function
P111(5) | 0 | Relay #5 control function
P111(6) | 0 | Relay #6 control function | Notes
P112(1) | 0 | Relay #1 ON setpoint | 1: Relays 1,2,4 and 5 are SPST. Relays
P112(2) | 0 | Relay #2 ON setpont | 3 and 6 are SPDT.
P112(3) | 3 | Relay #3 ON setpoint
P112(4) | 0 | Relay #4 ON setpoint
P112(5) | 0 | Relay #5 ON setpoint | 2: If relay assigned as level alarm (P111=1, relay is
P112(6) | 8 | Relay #6 ON setpoint | normally closed. If relay is assigned for pump
P113(1) | 0 | Relay #1 OFF setpoint | control, relay is normally open. Change P118
P113(2) | 0 | Relay #2 OFF setpont | from '2' to '3' to reverse logic.
P113(3) | 3.5 | Relay #3 OFF setpoint
P113(4) | 0 | Relay #4 OFF setpoint
P113(5) | 0 | Relay #5 OFF setpoint | 3: May simulate level by:
P113(6) | 7.5 | Relay #6 OFF setpoint | A: Enter PROGRAM mode.
P132 | 132 | Pump start delay (sec) | B: Go to P920, hit ENTER. Simulated
P133 | 10 | Pump power resumption delay (sec) | level will increase, then decrease by
P136 |  | Wall cling reduction | 1%/Second. Hit UP or DOWN arrows
P200 | 2 | mA output range (2= 4-20 mA) | to increase/decrease rate of change
P201 | 1 | mA output function (1= level) | up to 4%/second.
P202 | 1 | mA output allocation | C: Change P000 to -1 to activate pump and
P210 |  | 4 mA output value | control relays during simulation.
P211 |  | 20 mA output value | D: Go to RUN mode to exit simulation.
P219 | OFF | mA output failsafe
P801 |  | Range extension |  | For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Note on D16: Sets discrete input 1 to act as the source for a level reading override. Example: normally open high level switch closes to force reading to programmed high level (P065). Note: External 10-50 VDC required on switch circuit.
Note on D17: This default level (in engineering units) is substituted for the current reading when contact 1 is closed. Example: low level switch to turn off pumps in case of error in primary measurement.
Note on D18: The material level reported when a failsafe state is initiated. This also affects relay operation.
Note on D20: Sets the control algorithm to trip the relays. Level: Based on level setpoints ON and OFF. Relays normally closed. Change P118(1-6) from 2 to 3 to reverse logic In Bounds: When level enters the range between ON and OFF setpoints. Out of Bounds: When level exits the range between ON and OFF setpoints. Rate of Change: Based on Rate setpoints ON and OFF. Temperature: Based onTemperature setpoints ON and OFF. Loss of Echo (LOE): When echo is lost. Cable Fault: When the circuit to the transducer is open. Fixed Duty Assist: Fixed ON and OFF setpoints (no alternation). Allows multiple pumps to run simultaneously. Fixed Duty Backup: Fixed ON and OFF setpoints, Allows only one pump to run at a time. Ex: pump 1 turns off when pump 2 starts. Alternate Duty Assist: Alternates ON and OFF setpoints, and allows multiple pumps to run. MOST COMMON FUNCTION. Alternate Duty Backup: alternates ON and OFF setpoints, and allows only one pump to run at a time. Service Ratio Duty Assist: Pumps start based on ratios assigned in P122(1-6). Multiple pumps can run simultaneously. Service Ratio Duty Backup: Pumps start based on ratios assigned in P122(1-6). Only one pump to run at a time. First In First Out: Relay-on levels alternate. First relay to turn on is first relay to turn off. Flush Valve: Used to control a pump flushing device based on flush systems (P170 to P173). Communication: Based on input from external communications. See comms section of manual for info.
Note on D38: The minimum delay (in seconds) between pump starts. Use this to reduce a power surge from all pumps starting at the same time
Note on D39: The minimum delay (in seconds) before the first pump restarts after a power failure. This reduces the surge from the first pump starting immediately on power resumption.
Note on D40: This is the range (in engineering units) in which the relay setpoints are allowed to deviate. This randomly changes the ON and OFF setpoints within the range to reduce the material buildup on walls.
Note on D44: Level corresponding to 4 mA.
Note on D45: Level corresponding to 20 mA.
Note on D47: Material level may drop below 'zero level' by this amount before reporting Loss of Echo alarm

## well level tables (hidden)
Rate of Change
Slow | 1 | 2
Medium | 2
Fast | 3
Sensor type
ST-25 | 1 | 7
ST-50 | 2
ST-H | 100
XCT-8 | 101
XPS-10 | 102
XCT-12 | 103
XPS-15 | 104
XRS-5 | 112
mA input | 250
Units |  | 5
Meters | 1
Centimeters | 2 |  |  | 3
Millimeters | 3
Feet | 4
Inches | 5
Relay control function
Off |  |  | 0 | 1
Level |  |  | 1 | 1
In Bounds |  |  | 2 | 1
Out of Bounds |  |  | 3 | 1
Rate of Change |  |  | 4 | 1
Temperature |  |  | 5 | 1
Loss of Echos (LOE) |  |  | 6
Cable Fault |  |  | 7
Fixed duty assist |  |  | 50
Fixed duty backup |  |  | 51
Alternate duty assist |  |  | 52
Alternate duty backup |  |  | 53
Service ratio duty assist |  |  | 54
Service ratio duty backup |  |  | 55
First in, first out |  |  | 56
Flush Valve |  |  | 64
Communication |  |  | 65
P219, Milliamp output failsafe
mA output = failsafe level (P071) |  |  | OFF | 1
Last known value |  |  | HOLD
Empty mA output |  |  | LO
Span mA value |  |  | HI
P071, Failsafe material level
Level in units |  |  |  | 4
Level goes to maximum span |  |  | HI
Level goes to zero span (empty) |  |  | LO
Level remains at last reading |  |  | HOLD
P064, Auxiliary high or low level switch
Off. Not used |  |  | 0 | 1
On. High or low level switch wired to input 1 |  |  | 1
11

## Travelling screen data (hidden)
Rate of Change
Slow | 1 | 3
Medium | 2
Fast | 3
Sensor type
ST-25 | 1 | 7
ST-50 | 2 | 7
ST-H | 100
XCT-8 | 101
XPS-10 | 102
XCT-12 | 103
XPS-15 | 104
XRS-5 | 112
mA input | 250
Units
Meters | 1 | 5
Centimeters | 2
Millimeters | 3
Feet | 4
Inches | 5
Relay control function
Off |  |  | 0 | 2
Level |  |  | 1 | 2
In Bounds |  |  | 2 | 1
Out of Bounds |  |  | 3 | 1
Rate of Change |  |  | 4 | 1
Temperature |  |  | 5 | 1
Loss of Echos (LOE) |  |  | 6
Cable Fault |  |  | 7
Fixed duty assist |  |  | 50
Fixed duty backup |  |  | 51
Alternate duty assist |  |  | 52
Alternate duty backup |  |  | 53
Service ratio duty assist |  |  | 54
Service ratio duty backup |  |  | 55
First in, first out |  |  | 56
Flush valve |  |  | 64
Communication |  |  | 65
P219, Milliamp output failsafe
Ma output = failsafe level (P071) |  |  | OFF | 1
Last known value |  |  | HOLD
Empty mA output |  |  | LO
Span mA value |  |  | HI
P071, Failsafe material level
Level in units |  |  |  | 3
Level goes to maximum span |  |  | HI
Level goes to zero span (empty) |  |  | LO
Level remains at last reading |  |  | HOLD
P064, Auxiliary level switch
Off. Not used |  |  | 0 | 1
On. High or low level switch wired to input 1 |  |  | 1
Sensor type, input 2, P004 (2)
ST-25 | 1 | 6
ST-50 | 2
ST-H | 100
XCT-8 | 101
XPS-10 | 102
XCT-12 | 103
XPS-15 | 104
XRS-5 | 112
mA input | 250
Output allocation (P202)
Point 1 Level sensor 1 |  | 1 | 3
Point 2, level sensor 2 |  | 2
Point 3, differential |  | 3
Relay 1 source (P110)
Transducer 1 |  | 1 | 1
Transducer 2 |  | 2 | 1
Differential |  | 3 | 1
1
1
1
Relay 1 source (P110)
Transducer 1 |  | 1 | 3
Transducer 2 |  | 2 | 3
Differential |  | 3 | 3
1
1
2

## Travelling screen (hidden)
Siemens HydroRanger 200, Multiranger 200, Enviroranger Travelling Screen Control |  |  |  |  | Rate of change | Relay 1 source
Location:
Date:
[5 empty] | Sensor type, input 1 | Relay 2 source
Parameter | Value | Description
P001 | 4 | Operation (4= differential)
P002 | 1 | Material (1= liquid)
P003 (1,2) | 3 | Maximum process speed |  |  | Sensor type, input 2 | Relay 3 source
P004 (1) | 104 | Sensor type (input 1)
P004 (2) | 104 | Sensor type (input 2)
P005 | 5 | Units
P006 (1) |  | Distance, sensor face (1) to min level |  |  | Units | Relay 4 source
P006 (2) |  | Distance, sensor face (2) to min level
P007 (1) |  | Max. level above zero level, sensor (1)
P007 (2) |  | Max. level above zero level, sensor (2) |  |  | Failsafe material level | Relay 5 source
P064 | 0 | Reading override enable
P065 |  | Reading override value
P071 | LO | Failsafe material level | Notes: |  | Output allocation | Relay 6 source
P110 (1) | 1 | Level source, Relay #1 | 1: Relays 1,2,4 and 5 are SPST. Relays
P110 (2) | 1 | Level source, Relay #2 | 3 and 6 are SPDT.
P110 (3) | 1 | Level source, Relay #3
P110 (4) | 1 | Level source, Relay #4 | 2: May simulate level by: |  | Output failsafe
P110 (5) | 1 | Level source, Relay #5 | A: Enter PROGRAM mode.
P110 (6) | 1 | Level source, Relay #6 | B: Go to P920, hit ENTER. Simulated |  |  | Relay 1 control function
P111(1) | 1 | Relay #1 control function | level will increase, then decrease by
P111(2) | 1 | Relay #2 control function | 1%/Second. Hit UP or DOWN arrows |  | Auxiliary high or low level switch
P111(3) | 0 | Relay #3 control function | to increase/decrease rate of change
P111(4) | 0 | Relay #4 control function | up to 4%/second. |  |  | Relay 2 control function
P111(5) | 0 | Relay #5 control function | C: Change P000 to -1 to activate pump and
P111(6) | 0 | Relay #6 control function | control relays during simulation.
P112(1) |  | Relay #1 ON setpoint | D: Go to RUN mode to exit simulation.
P112(2) |  | Relay #2 ON setpont |  |  |  | Relay 3 control function
P112(3) |  | Relay #3 ON setpoint | 3: If relay assigned as level alarm (P111=1, relay is
P112(4) |  | Relay #4 ON setpoint | normally closed. If relay is assigned for pump
P112(5) |  | Relay #5 ON setpoint | control, relay is normally open. Change P118
P112(6) |  | Relay #6 ON setpoint | from '2' to '3' to reverse logic. |  |  | Relay 4 control function
P113(1) |  | Relay #1 OFF setpoint
P113(2) |  | Relay #2 OFF setpont
P113(3) |  | Relay #3 OFF setpoint
P113(4) |  | Relay #4 OFF setpoint |  |  |  | Relay 5 control function
P113(5) |  | Relay #5 OFF setpoint
P113(6) |  | Relay #6 OFF setpoint
P132 |  | Pump start delay (sec)
P133 |  | Pump power resumption delay (sec) |  |  |  | Relay 6 control function
P136 |  | Wall cling reduction
P200 | 2 | mA output range (2= 4-20 mA)
P201 | 1 | mA output function (1= differential)
P202 | 3 | mA output allocation
P210 |  | 4 mA output value
P211 |  | 20 mA output value
P219 | OFF | mA output failsafe
For more option, refer to instruction manual,
or call Siemens Service @ 800-333-7421.
Note on D11: Sensor type for input 2 must match sensor type input 1, or transducer for input 1 and milliamp input for transducer 2.
Note on D17: Sets discrete input 1 to act as the source for a level reading override. Example: normally open high or low level switch closes to force reading to programmed programmed level (P065). Note: External 10-50 VDC required on switch circuit.
Note on D18: This level (in engineering units) is substituted for the current reading when contact 1 is closed. Note: substitutes level 1 only, not differential.
Note on D19: The material level reported when a failsafe state is initiated. This also affects relay operation.
Note on D20: Sets the level source on which each of the relays is assigned to a measurement point. Example: Relays 1-3 may be assigned to differential level (value 3), while relays 4-6 may be assigned to level from sensor 1 (value 1).
Note on D26: Sets the control algorithm to trip the relays. Level: Based on level setpoints ON and OFF. In Bounds: When level enters the range between ON and OFF setpoints. Out of Bounds: When level exits the range between ON and OFF setpoints. Rate of Change: Based on Rate setpoints ON and OFF. Temperature: Based onTemperature setpoints ON and OFF. Loss of Echo (LOE): When echo is lost. Cable Fault: When the circuit to the transducer is open. Fixed Duty Assist: At fixed ON and OFF setpoints, and alows multiple pumps to run. Fixed Duty Backup: At fixed ON and OFF setpoints, and allows only one pump to run. Alternate Duty Assist: At rotating ON and OFF setpoints, and allows multiple pumps to run. Alternate Duty Backup: At rotating ON and OFF setpoints, and allows only one pump to run. Service Ratio Duty Assist: On service ratio ON and OFF setpoints and allows multiple pumps to run. Service Ratio Duty Backup: On service ratio ON and OFF setpoints and allows only one pump to run. First In First Out: As alternate duty assist, resets the relay from staggered OFF setpoints. Flush Valve: Used to control a pump flushing device based on flush systems (P170 to P173). Communication: Based on input from external communications. See comms section of manual for info.
Note on D44: The minimum delay (in seconds) between pump starts. Use this to reduce a power surge from all pumps starting at the same time
Note on D45: The minimum delay (in seconds) before the first pump restarts after a power failure. This reduces the surge from the first pump starting immediately on power resumption.
Note on D46: This is the range (in engineering units) in which the relay setpoints are allowed to deviate. This randomly changes the ON and OFF setpoints within the range to reduce the material buildup on walls.
Note on D49: Sets the input source from which the mA output is calculated.
Note on D50: Level, or differential corresponding to 4 mA.
Note on D51: Level, or differential corresponding to 20 mA.

## Hydro HMI tables (hidden)
Rate of Change |  |  | Relay source
Slow | Slow | 2 | Transducer 1 | Transducer 1 | 1
Medium | Medium |  | Transducer 2 | Transducer 2 | 2
Fast | Fast |  | Difference or Average | Difference or Average | 3
1
1
1
Sensor type
ST-25 | ST-25 | 7 | Relay control function
ST-50 | ST-50 |  | Off | Off | 1
ST-H | ST-H |  | Level | Level | 1
XCT-8 | XCT-8 |  | Low Low Alarm | Low Low Alarm | 1
XPS-10 | XPS-10 |  | Low Alarm | Low Alarm | 1
XCT-12 | XCT-12 |  | High Alarm | High Alarm | 1
XPS-15 | XPS-15 |  | High High Alarm | High High Alarm | 1
XRS-5 | XRS-5 |  | In Bounds | In Bounds
mA input | mA input |  | Out of Bounds | Out of Bounds
|  |  | Rate of Change | Rate of Change
|  |  | Temperature | Temperature
Units |  |  | Loss of Echos (LOE) | Loss of Echos (LOE)
Meters | Meters | 5 | Cable Fault | Cable Fault
Centimeters | Centimeters |  | Fixed duty assist | Fixed duty assist
Millimeters | Millimeters |  | Fixed duty backup | Fixed duty backup
Feet | Feet |  | Alternate duty assist | Alternate duty assist
Inches | Inches |  | Alternate duty backup | Alternate duty backup
|  |  | Service ratio duty assist | Service ratio duty assist
|  |  | Service ratio duty backup | Service ratio duty backup
|  |  | First in, first out | First in, first out
|  |  | Flush valve | Flush valve
|  |  | Communication | Communication

## Hydroranger_HMI Differential
Siemens Hydroranger-200_HMI Differential Level
| Location: | [6 empty] | Units
| Date: | 2026-07-29
| Go to Setup | [6 empty] | Response Rate
Sensor
| Units (universal for both channels) | Inches
| Level Selector | 1 | 2 | 3 |  |  |  | Sensor Type
| Sensor Mode | Level | Level | Dual Point Difference
| Transducer Selector | 1 | 2 | NA
| Transducer | XPS-15 | XPS-15 | NA
| Material | Liquid | Liquid | NA
| Calibration | [6 empty] | Relay 1 Source
| Level Selector | 1 | 2 | 3
Span
| Transducer Selector | 1 | 2 | NA |  |  |  | Relay 2 Source
| Empty |  |  | NA
| Sensor Offset |  |  | NA
| Blanking Distance (11.8" default) |  |  | NA |  |  |  | Relay 3 Source
| Range Extension (20% of Empty default) |  |  | NA
Rate
| Response Rate (Slow, Med, Fast) | Medium | Medium | NA |  |  |  | Relay 4 Source
Current Output
| MA Output Selector (AO1 or AO2) | 1 | 2 | NA
| MA Output Range | 4 to 20 ma | 4 to 20 ma | NA |  |  |  | Relay 5 Source
| Curr. Out Func. | Level | Level | NA
| MA output Allocation (Point 1, 2, 3) |  |  | NA
| 4 MA Setpoint |  |  | NA |  |  |  | Relay 6 Source
| 20 MA Setpoint |  |  | NA
Relays/Basic Setup
| Relay selector | 1 | 2 | 3 | 4 | 5 | 6
| Level Source (xdcr 1, 2, difference) | Transducer 1 | Transducer 2 | Difference or Average | Transducer 1 | Transducer 1 | Transducer 1
| Preset Applications | Off | Off | Off | Off | Off | Off | Relay 1 Control Function
| Relay Function | Off | Off | Off | Off | Off | Off
On Setpoint
| Off Setpoint | [6 empty] | Relay 2 Control Function
Relay Logic (Positive, Negative)
Relay 3 Control Function
Relay 4 Control Function
Relay 5 Control Function
Relay 6 Control Function
Note on B28: Select as follows: 1 if level from sensor 1 2 if level from sensor 2 3 if differential level
Note on B30: If milliamp allocation set to '3' (differential), this value sets at what differential level 20 mA occurs. Normally same as Calibration/Span.
Note on B38: Positive If relay is programmed for alarm, relay is normally energized, and de-energizes on alarm. If relay is programmed for pump control, relay is normally de-energized, and energizes upon exceeding setpoint. Negative Opposite of above

## LT500 pump control files (hidden)
Units |  |  | Number of relays
Meters | 5 |  | 0 | [12 empty] | Relay 1 | Relay 2 | Relay 3 | Relay 4 | Relay 5 | Relay 6 | 4
Centimeters |  |  | 1 | Pump 1 | [5 empty] | 1 | [6 empty] | Relay 2 | Relay 3 | Relay 4 | Relay 5 | Relay 6
Millimeters |  |  | 2 | Pump 1 | Pump 2 |  |  |  |  | 1 | 2 | [6 empty] | Relay 3 | Relay 4 | Relay 5 | Relay 6
Feet |  |  | 3 | Pump 1 | Pump 2 | Pump 3 |  |  |  | 1 | 2 | 3 | [6 empty] | Relay 4 | Relay 5 | Relay 6
Inches |  |  | 4 | Pump 1 | Pump 2 | Pump 3 | Pump 4 |  |  | 1 | 2 | 3 | 4 | [6 empty] | Relay 5 | Relay 6
|  |  | 5 | Pump 1 | Pump 2 | Pump 3 | Pump 4 | Pump 5 |  | 1 | 2 | 3 | 4 | 5 | [6 empty] | Relay 6
|  |  | 6 | Pump 1 | Pump 2 | Pump 3 | Pump 4 | Pump 5 | Pump 6 | 1 | 2 | 3 | 4 | 5 | 6
Sensor type
LR110 | 4
LR120
Probe LU240
4-20 mA
Response rate
Slow | 1
Medium
Fast
Pump control mode | 1
Alternate duty assist | 1
Alternate duty backup
Service ratio duty assist
Service ratio duty backup
Fixed duty assist
Fixed duty backup | 1
View 1 |  | View 2
1 value | 3 | 1 value | 6: 7
1 Value and horizontal bargraph |  | 1 Value and horizontal bargraph
1 value, icons and vertical bargraph |  | 1 value, icons and vertical bargraph
1 value and trend chart |  | 1 value and trend chart
3 values and horizontal bargraph |  | 3 values and horizontal bargraph
3 values, icons and vertical bargraph |  | 3 values, icons and vertical bargraph
6 values |  | 6 values
1st value |  | 1st value
Level (point 1) | 1 | Level (point 1) | 6: 2
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference |  | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
2nd value |  | 2nd value
Level (point 1) | 8 | Level (point 1) | 6: 3
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference | 2 | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
None |  | None
3rd value | 1 | 3rd value
Level (point 1) | 8 | Level (point 1) | 6: 6
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference |  | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
None |  | None
4th value |  | 4th value
Level (point 1) | 8 | Level (point 1) | 6: 2
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference |  | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
None |  | None
5th value |  | 5th value
Level (point 1) | 8 | Level (point 1) | 6: 4
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference |  | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
None |  | None
6th value |  | 6th value
Level (point 1) | 8 | Level (point 1) | 6: 7
Level (point 2) |  | Level (point 2)
Distance (point 1) |  | Distance (point 1)
Distance (point 2) |  | Distance (point 2)
Level difference |  | Level difference
Current output 1 |  | Current output 1
Current output 2 |  | Current output 2
None |  | None

## Keystroke instructions
Before Programming, erase
existing configuration
1. In program model, access P999.
2. Press C (clear) button.
3. Press Enter button.
4. C.ALL displays until the reset is
complete